## Sheet1
No. | Date | QC supervisor | Supervisor | Inspector Name | Vendor Name | Vendor Code | Division | Item | PO | Lot | PO cases | PO Eaches | Rejection reason ( Add with 100% in description) | Defect Category | Disposition | Key Root Cause | Correction plan | Comments
1 | 2022-01-07 | Thailand | Chutima Ruensamran |  | COBES INDUSTRIES (MYANMAR) COMPANY LIMITED | 143305 | 21 | DYNJP2001S | 4514518738 | 30521120059 |  |  | The Manufacturing Date on pouch labelling and Mfg. date on carton does not match | Labeling | Accept by Deviation
2 | 2022-01-07 | India | K. Dinakaran |  | NEW WORLD IMPORTS | 147324 | 81 | MSC095014 | 4514507532 | 2135920/2136020/2136021/2136121/2136122/2136222/2136322/2136323/2136423/2136523/2136524/2200125 |  |  | 1.Lot Mix up --- Found lot#2136222 Cans in Lot Case of 2136322.>>01 Case 2.Lot Mix up --- Found lot#2136122 Cans in Lot Case of 2136222.>>01 Case | Visual | Rework |  | As Oversight Mistake, Cans were mixed up during the Shift Change times leads Ineffective control over Line Clearance. Workers were Trained to Control of Line Clearance during the Closing of the Shift.
3 | 2022-01-10 | Hefei | Echo Liu |  | SINO PROTECTION HOLDINGS CO LTD | 105110 | ISP-EU | TB29433CE | SPT11191JB | 22AAS049 |  |  | 100% sample dimensions were out of tolerance. The width of drape should be 76.5”-77.5”. The actual products were all 75” and 75.5”. | Dimensional | Accept by Deviation |  | Fix machine
4 | 2022-01-12 | Shanghai 2 | Joy Zhou |  | SUZHOU MECCAN IMP AND EXP CORP | 12911 | 81 | TFX00286204 | N/A | 91121120001 |  |  | No label on case carton，100% | Labeling | Rework |  | Re-explain PIC'S requirments for factory
5 | 2022-01-16 | India | K. Dinakaran |  | RR GLOBAL TRADING FZE | 152936 | 51 | NON50240F | 4514416847 | 93621120001 |  |  | wrist opening measurement out of tolerance | Dimensional | Accept by Deviation |  | Vendor Made an Oversight mistake as Missed verifying the Updated Spec Sheet during production. In continuation, Vendor updated the spec sheet in their records and Made New Patterns as per the Updated Spec Sheet.
6 | 2022-01-16 | India | K. Dinakaran |  | RR GLOBAL TRADING FZE | 152936 | 51 | NON50240XLF | 4514416847 | 93621120001 |  |  | wrist opening measurement out of tolerance | Dimensional | Accept by Deviation |  | Vendor Made an Oversight mistake as Missed verifying the Updated Spec Sheet during production. In continuation, Vendor updated the spec sheet in their records and Made New Patterns as per the Updated Spec Sheet.
7 | 2022-01-16 | India | K. Dinakaran |  | RR GLOBAL TRADING FZE | 152936 | 51 | NON50240F | 4514416837 | 93621120002 |  |  | wrist opening measurement out of tolerance, 100% | Dimensional | Accept by Deviation |  | Vendor Made an Oversight mistake as Missed verifying the Updated Spec Sheet during production. In continuation, Vendor updated the spec sheet in their records and Made New Patterns as per the Updated Spec Sheet.
8 | 2022-01-16 | India | K. Dinakaran |  | RR GLOBAL TRADING FZE | 152936 | 51 | NON50240XLF | 4514416837 | 93621120002 |  |  | wrist opening measurement out of tolerance | Dimensional | Accept by Deviation |  | Vendor Made an Oversight mistake as Missed verifying the Updated Spec Sheet during production. In continuation, Vendor updated the spec sheet in their records and Made New Patterns as per the Updated Spec Sheet.
9 | 2022-01-18 | Hefei | Echo Liu |  | PRINCIPLE & WILL BIOTECH PINGHU CO | 148292 | 22 | SPT-2202CS | 4600064584 | 29222010010PH |  |  | SPT-2202CS with "Neck To Reinforced Top" out of tolerance,spec is 5” Min, but actually all with 4.5” | Dimensional | Accept by Deviation |  | Retrain worker
10 | 2022-01-18 | Shanghai 2 | Peter Zhang |  | HANGZHOU FORMED MEDICAL DEVICES CO | 147635 | 29 | DYNJNASOF28 DYNJNASOF32 | 4514560915 | 63522010001 |  |  | All/ Item number description incorrect on the IM Required: DYNJNASOF28,DYNJNASOF32 Actual: DDYNJNASOF28,DDYNJNASOF32 | Visual | Rework |  | Re-explain PIC'S requirments for factory
11 | 2022-01-27 | India | K. Dinakaran |  | MEDIBLUE HEALTHCARE PVT LTD | 152280 | 51 | NON27203M | 4514463989/4514463990 | 28022010008 |  |  | Size Mix up | Visual | Rework |  | Stoped stitching similar styles like L,M,XL in single sewing line to avoid mix-up`s.
12 | 2022-02-06 | India | K. Dinakaran |  | JIEHONG MEDICAL PRODUCTS BD CO LTD | 148895 | 51 | NON27213M | 4514598738 | 89522010002 |  |  | Found two different color variation on the panel, Holes and also Loose thread found on the product . For this reason this lot has been rejected. | Visual | Rework |  | Due to the Oversight Mistake, Workers skipped to verify the said Defect. Training was given to workers how to check this defect.
13 | 2022-02-11 | Shanghai1 | Joy Zhou |  | SHANGHAI MINHUA CO | 117076 | 81 | MDS630XSQ | 4514389284 | 7621120003 |  |  | 100% printing wrong ,The text in the drawing does not match exactly with actual text, drawing text is "behind thigh", actually is "behind knee" | Visual | Accept by updating PIC |  | Minhua Update Drawings
14 | 2022-02-13 | Shanghai1 | Joy Zhou |  | BYTECH DONGTAI CO LTD | 146587 | 40 | VEN1235 | 4514585108 | BY202585108 |  |  | 100% actual lot number on pack is BY201560899, different from the lot on carton and outer bag. | Labeling | Accept by Deviation
15 | 2022-02-13 | Shanghai1 | Joy Zhou |  | BYTECH DONGTAI CO LTD | 146587 | 40 | VEN1230 | 4514585108 | BY202585108 |  |  | 100% actual lot number on pack is BY201560899, different from the lot on carton and outer bag. | Labeling | Accept by Deviation
16 | 2022-02-18 | Shanghai 2 | Peter Zhang |  | HANGZHOU GSPMED MEDICAL | 140085 | 81 | MSC6800 | 4514573622 | 8521120001 |  |  | For 304cases, the old version of case label and pouch were used, but the new version of inner box was used. For 368cases, the new vsrsion of cases label, inner box and pouch were used. Old versions of labels in stock were used up. New versions of labels will used for future orders. 10ea, length of Cohesive Bandage was less than 3.5m. Actually found: 3.49, 3.42, 3.25, 3,43, 3,39, 3,30, 3,42, 3,36, 3,30, 3.46. | Labeling | Accept by Deviation |  | Re-explain PIC'S requirments for factory
17 | 2022-02-18 | Shanghai 2 | Peter Zhang |  | HANGZHOU GSPMED MEDICAL | 140085 | 81 | MSC6800XL | 4514573622 | 8521120001 |  |  | For all 96cases, the old version of case label and pouch were used, but the new version of inner box was used. Old versions of case label and pouch in stock will still be used for PO#4514731519. | Labeling | Accept by Deviation |  | Re-explain PIC'S requirments for factory
18 | 2022-02-18 | Shanghai 2 | Peter Zhang |  | HANGZHOU GSPMED MEDICAL | 140085 | 81 | MSC6802 | 4514573622 | 8521120001 |  |  | For all 288cases, the old version of case label and pouch were used, but the new version of inner box was used. Old versions of case label and pouch in stock will still be used for PO#4514731519. | Labeling | Accept by Deviation |  | Re-explain PIC'S requirments for factory
19 | 2022-02-19 | Hefei | Echo Liu |  | SINO PROTECTION HOLDINGS CO LTD | 105110 | 22 | 22706P | 9000837381 | 11022020178 |  |  | SINO-HEFEI item 22706P SINO use the different material the body is PFG 85g Blue, the neck binding is SPP 35g Dark blue. While the GPI states : the neck binding material is same as gown body PFG 85g Blue. | Visual | Rework |  | Update Spec.
20 | 2022-03-04 | Hubei | Cherry Cheng |  | TOPWIDE HUBEI MEDICAL PRODUCTS | 137125 | 51 | NONSW025XL | 4514598727/4514598745 | 12522030040 |  |  | Product description on outer case labeling was incorrect. It should be hook &loop instead of snap | Labeling | Rework |  | Medline update labeling to vendor
21 | 2022-03-07 | Shanghai 2 | Peter Zhang |  | BESMED HEALTH BUSINESS CORP | 137015 | 35 | HCSU4486 | 4514622870 | 8522020325 |  |  | Barcode scanned incorrect | Labeling | Rework |  | Provide barcode training for vendor
22 | 2022-03-14 | NOC | Shawn Wei |  | DALIAN SUNDLY HOME PRODUCTS CO LTD | 152138 | 82 | MSC282045 | 4514588830 | 13822020001 |  |  | The core length dimension was short. Our spec is 68+/-2.5. Sundly makes it 65. | Dimensional | Accept by Deviation |  | Align with vendor about our Spec for their sales and manufacturing
23 | 2022-03-15 | Shanghai1 | Joy Zhou |  | ZHANGJIAGANG WEIFENGRUN PLASTIC CO | 145419 | 40 | CLE2000HDPE | 4514330547 | WF214330547 |  |  | 100% incorrect lot# outer/inner case/package labeling | Labeling | Accept by Deviation |  | correct the lot# from the new order
24 | 2022-03-15 | Shanghai1 | Joy Zhou |  | JIANGSU XINBEI ELECTRICAL APPLICANC | 142500 | 81 | MDS67186 | 202151 | 50022020003 |  |  | 100% Incorrect PO# and Lot#. | Visual | Accept by Deviation
25 | 2022-03-16 | Shanghai 2 | Peter Zhang |  | RUIDE POLYURETHANE FOAM ZHEJIANG CO | 154009 | 15 | NON081449 | 4514612533 | 2112-2109 |  |  | Dimensions out of tolerance | Dimensional | Accept by updating PIC |  | Re-explain PIC'S requirments for factory
26 | 2022-03-17 | NOC | Shawn Wei |  | HENAN HENGXIANG MEDICAL SUPPLIES CO | 120480 | 50 | MDTBB3B17R | 4514712313 | HEH22C |  |  | incorrect bale packaging method, not sewing four sides of poly propylene sheet securely | Visual | Rework |  | Follow the diagram spec. and method confirmed with QA for packaging.
27 | 2022-03-17 | NOC | Shawn Wei |  | HENAN HENGXIANG MEDICAL SUPPLIES CO | 120480 | 50 | MDTBB4B20R | 4514712313 | HEH22C |  |  | incorrect bale packaging method, not sewing four sides of poly propylene sheet securely | Visual | Rework |  | Follow the diagram spec. and method confirmed with QA for packaging.
28 | 2022-03-18 | NOC | Shawn Wei |  | HONGRAY (USA) MEDICAL PRODUCTS INC | 130324 | ISP-Japan | CUR8234J | HMVJ203004 | BIN22030010J |  |  | 100% Box color is lighter than approved artwork. | Labeling | Rework |  | Compare the signed box color and artwork color for every lot.
29 | 2022-03-18 | India | K. Dinakaran |  | MEDIBLUE HEALTHCARE PVT LTD | 152280 | 51 | NON27209S | 4514675041/4514742639 | 28022030013 |  |  | Product Mix up | Visual | Rework |  | Lack of Awarness to new workers:- Awarness training given to new workers to check the lables before packing.
30 | 2022-03-21 | Shanghai 2 | Peter Zhang |  | NINGBO LIGHTPAK IMP AND | 120427 | 20 | 500492 | 9000838494 | 42722040002 |  |  | 1. Actual found the side pocket was on the product right instead of left stated on specification; The vendor explains that this does not affect the use of the function | Visual | Rework |  | Provide training for factory, they must follow PIC's requirements
31 | 2022-03-22 | Dongguan | Tim Li |  | JIE GAO PLASTICS PRODUCTS CO LTD | 120930 | 29 | 58008 | 9000838334 | 93022033414U |  |  | 100% Wrong LOT No(93022033414) on the carton. | Labeling | Rework |  | 1. Retrain QC for all requirements check. 2. Require QC to refer to first article sample during each inspection.
32 | 2022-03-25 | Shanghai1 | Joy Zhou |  | ZHANGJIAGANG WEIFENGRUN PLASTIC CO | 145419 | 40 | CLE2000HDPE | 4514759916 | WF214759916 |  |  | 100% with incorrect lot number per PIC-00569 on package | Labeling | Accept by Deviation
33 | 2022-03-25 | Shanghai1 | Joy Zhou |  | SHEN WEI USA INC | 102595 | 40 | CR911XL | 4514773474 | AZ107452634 |  |  | 100% with wrong PO & lot numbers printed on packaging.due to we released order but cancelled before, so there are some stock with the old PO & Lot printed on package. we agreed vendor ship out old stock with the new order. | Labeling | Accept by Deviation
34 | 2022-03-25 | Shanghai1 | Joy Zhou |  | SHEN WEI USA INC | 102595 | 40 | VEN6145W | 4514788133 | AZ111343019 |  |  | 100% with wrong PO & lot numbers printed on packaging.due to we released order but cancelled before, so there are some stock with the old PO & Lot printed on package. we agreed vendor ship out old stock with the new order. | Labeling | Accept by Deviation
35 | 2022-03-25 | Dongguan | Tim Li |  | MASTER AND FRANK ENTERPRISE | 105450 | 22 | SPT-2550A | 9000838035 | 45022030171HJ |  |  | 32pcs stamp position and stamp arrow direction can't match with approved drawing.100% | Visual | Rework |  | 1.Retrain the line leader and QC on related inspection specification. 2.Require line leader to check first article sample against product specification before production. 3.Require QC to inspect product with product specification carefully during first article/in-process/final inspection.
36 | 2022-03-28 | Shanghai 2 | Peter Zhang |  | ZHEJIANG WELLONG MEDICAL TECHNOLOGY | 148639 | 20 | MSC2104 | 201820 | 63921120900 |  |  | length out of the Tolerance“125 ±3 mm” in PIC-00286 (as attached picture 001 ,actual length A of adhesive film is 109mm, length B of stripping paper is 120mm) | Dimensional | Accept by updating PIC |  | Re-explain PIC'S requirments for factory
37 | 2022-03-28 | Shanghai 2 | Peter Zhang |  | ZHEJIANG WELLONG MEDICAL TECHNOLOGY | 148639 | 20 | MSC2002 | 201820 | 63921120900 |  |  | length out of the Tolerance 76 ±3 mm in PIC-00286,actual length of adhesive film is 65 ±3mm as attached picture 004 Dimension c; | Dimensional | Accept by updating PIC |  | Re-explain PIC'S requirments for factory
38 | 2022-03-29 | Shanghai1 | Joy Zhou |  | WU XI LONG TENG TEXTILE CO LTD | 123001 | 50 | MDT211434PB | 4514833753 | WTC22C |  |  | 100% Incorrect packaging./incorrect product.The actual packaging is 250ea/cs, 5 bag/cs,50ea/bag.it should be 500ea/cs,50ea/box,10box/cs. The sewn in seam was on the outside which should be on the inside of the cap | Labeling | Accept by Deviation |  | input the correction plan
39 | 2022-03-29 | Shanghai1 | Joy Zhou |  | WU XI LONG TENG TEXTILE CO LTD | 123001 | 50 | MDT211434FPB | 4514833753 | WTC22C |  |  | 100% Incorrect packaging:The bottom of the hats were not folded. | Visual | Accept by Deviation
40 | 2022-03-30 | Thailand | Chutima Ruensamran |  | SESMED (CAMBODIA) HEALTHCARE | 143208 | 21 | DYNJP2202S | 4514706395 | 20822030013K |  |  | 200x Incorrect Stamp (DYNJP2202), the correct should be DYNJP2202S.100% | Visual | Rework
41 | 2022-04-01 | Hefei | Echo Liu |  | SINO PROTECTION HOLDINGS CO LTD | 105110 | 22 | 129616 | 9000839777 | 11022070124 |  |  | The description of product name on carton is incorrect,rate 100% | Labeling | Rework |  | Update artwork
42 | 2022-04-01 | Shanghai 2 | Joy Zhou |  | SHANGHAI MINHUA CO | 117076 | 10 | MDS9911HPS | 4514631534 | 20220223MST |  |  | The product printing is incorrect, 100% | Visual | Rework |  | Provide training for vendor, and strengthen process checking
43 | 2022-04-06 | Shanghai1 | Joy Zhou |  | SHEN WEI USA INC | 102595 | 40 | VEN6145W | 4514788135 | AZ111403340 |  |  | 100% with wrong PO & lot numbers printed on packaging.due to we released order but cancelled before, so there are some stock with the old PO & Lot printed on package. we agreed vendor ship out old stock with the new order. | Labeling | Accept by Deviation
44 | 2022-04-07 | NOC | Shawn Wei |  | INTCO MEDICAL HK CO LIMITED | 125123 | 40 | MDS2584 | 4514874230 | ZIM204874230 |  |  | no film at the dispenser box opening | visual | Rework |  | strickly follow our specification and double check
45 | 2022-04-08 | Shanghai1 | Joy Zhou |  | SHEN WEI USA INC | 102595 | 40 | CLE2000HDPE | 4514788136 | AZ111403341 |  |  | 100% with incorrect lot number per PIC-00569 on package | Labeling | Accept by Deviation
46 | 2022-04-08 | Dongguan | Joy Zhou |  | YANGZHOU PERFECT DAILY | 123460 | ISP-Japan | GEMJ3172 | FNP011422A | FNP22030014J |  |  | 50EA folding method were wrong.100% | Visual | Rework |  | 1.Retrain the line leader and QC on related inspection specification. 2.Require line leader to check first article sample against product specification before production. 3.Require QC to inspect product with product specification carefully during first article/in-process/final inspection.
47 | 2022-04-12 | NOC | Shawn Wei |  | HENAN HENGXIANG MEDICAL SUPPLIES CO | 120480 | 50 | MDTR1082043BB | 4514932833 | HEH22C |  |  | Bale insert with wrong contract PO#0000202026 | Labeling | Rework |  | Print the formal PO# not contract PO for div.50.
48 | 2022-04-14 | Shanghai1 | Joy Zhou |  | SHEN WEI USA INC | 102595 | 40 | VEN6145W | 4514788136 | AZ111403341 |  |  | 100% with wrong PO & lot numbers printed on packaging.due to we released order but cancelled before, so there are some stock with the old PO & Lot printed on package. we agreed vendor ship out old stock with the new order. | Labeling | Accept by Deviation
49 | 2022-04-15 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | ISP-EU | NONE27374P | XSP030722B | 22DASP175 |  |  | Nose piece length is smaller than SPEC. SPEC: 9.5cm to 11cm Defect: 9cm | Dimensional | Rework |  | Vendor corrected next production
50 | 2022-04-19 | Shanghai1 | Joy Zhou |  | SUZHOU JINGLE POLYMER MEDICAL | 118486 | ISP-Japan | DYND371603J | SUJ010722W | SUJ22040013J |  |  | 100% paper scrapes on the outside of pouch | Visual | Accept by Deviation
51 | 2022-04-23 | Hefei | Echo Liu |  | SINO PROTECTION HOLDINGS CO LTD | 105110 | ISP-EU | NS-8510 | SPT03102ES | 22DAS108 |  |  | The dimension of wrap is incorrect. | Dimensional | Accept by Deviation |  | first sample confirmation and FQC,IPQC
52 | 2022-04-28 | Shanghai1 | Joy Zhou |  | BYTECH DONGTAI CO LTD | 146587 | 40 | CLE204 | 4514851419 | BY204851419 |  |  | 100% with wrong PO & lot numbers printed on packaging.due to we released order but cancelled before, so there are some stock with the old PO & Lot printed on package. we agreed vendor ship out old stock with the new order. | Labeling | Accept by Deviation
53 | 2022-05-07 | Shanghai1 | Joy Zhou |  | WUHU LIXIN CLEANING PRODUCTS CO | 151240 | 22 | 129258 | 9000838966 | 24022050001 |  |  | 100%Missing item description on case label | Labeling | Accept by Deviation
54 | 2022-05-09 | Shanghai 2 | Peter Zhang |  | IMMANUEL INDUSTRIAL COMPANY LIM | 142121 | 81 | NON024230 | 4514919528 | 12122032401 |  |  | Weight out of tolerance Actual :10 pcs :13.3-13.5g | Dimensional | Accept by updating PIC |  | Re-explain PIC'S requirments for factory
55 | 2022-05-12 | NOC | Shawn Wei |  | HONGRAY (USA) MEDICAL PRODUCTS INC | 130324 | 40 | VEN1240V1 | 4514949128 | ME204880718 |  |  | Incorrect quantity per labelled pack | Labeling | Rework |  | Follow PIC spec. on each level of package unit for new items.
56 | 2022-05-13 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | ISP-EU | CRIE4040C | XSP030722A | 22DASP175 |  |  | The IFU and case/bag labels were not identical. It using old revision CRIE4040C_LAB210362_361 but with new revision IFU which with UKCA and CH REP, QA confirmed all level labelling must be identical, if you have UKCA and CH rep on one label, you must find them on the other levels. | Labeling | Rework |  | Vendor corrected next production
57 | 2022-05-18 | Shanghai1 | Joy Zhou |  | WUXI NEW BERRY INTERNATIONAL TRADIN | 146421 | 50 | MDT217649 | 4514937024 | 42122050004 |  |  | 100% item# is incorrect on the tag of product | Visual | Accept by Deviation
58 | 2022-05-23 | Shanghai1 | Joy Zhou |  | JIANERKANG MEDICAL CO LTD | 109260 | 81 | MDT2168206XR | 4514965658 | 26022050018 |  |  | 100%Missing yellow thread on seams and missing X-ray filament | Visual | Accept by Deviation
59 | 2022-05-23 | Dongguan | Joy Zhou |  | YANGZHOU PERFECT DAILY | 123460 | 50 | KAREBAG01 | WNR0320221 | WNR22W22 |  |  | 8CS/32EA Actual LOT# format on the case and IFU don't comply with spec Actual lot# : WNR22W22 Labeling spec: 76122W22 100% | Labeling | Rework |  | 1. Make sure the updated artwork being issued to production line before production. 2. Request Medline to send updated artwork to two vendor contacts, add one back-up to avoid ignorance.
60 | 2022-05-24 | Shanghai1 | Joy Zhou |  | BYTECH DONGTAI CO LTD | 146587 | 40 | CLE202 | 4514948077 | BY113353803,BY113353804, BY113353807 |  |  | 100% with wrong PO & lot numbers printed on packaging.due to we released order but cancelled before, so there are some stock with the old PO & Lot printed on package. we agreed vendor ship out old stock with the new order. | Labeling | Accept by Deviation
61 | 2022-06-01 | Shanghai1 | Joy Zhou |  | BYTECH DONGTAI CO LTD | 146587 | 40 | VEN1230V1 | 4514948083 | BY205948083 |  |  | 100% with actual lot number on pack is BY012646634, should be BY205948083, the lot on carton is BY205948083. | Labeling | Accept by Deviation
62 | 2022-06-02 | Dongguan | Joy Zhou |  | SUPREME SOURCE & SUPPLY CO | 149771 | 81 | NON200012 | 4514406916 | 77121120001 |  |  | 8CS/Missing label on the inner box 100% | Labeling | Rework |  | 1. Make sure the updated artwork being issued to production line before production. 2. Request Medline to send updated artwork to two vendor contacts, add one back-up to avoid ignorance.
63 | 2022-06-02 | Dongguan | Joy Zhou |  | SUPREME SOURCE & SUPPLY CO | 149771 | 81 | NON200018 | 4514406916 | 77121120001 |  |  | 8CS/Missing label on the inner box 100% | Labeling | Rework
64 | 2022-06-02 | Dongguan | Joy Zhou |  | SUPREME SOURCE & SUPPLY CO | 149771 | 81 | NONPF787 | 4514406916 | 77121120001 |  |  | 8CS/Missing label on the inner box/100% | Labeling | Rework
65 | 2022-06-07 | Hefei | Echo Liu |  | SINO PROTECTION HOLDINGS CO LTD | 105110 | 21 | DYNJP4109 | 4515060406 | 11022070080 |  |  | labeling information of carton and pouch does not match, Product of Saudi Arabla print on carton, but made in china print on pouch. | Labeling | Rework |  | Update artwork
66 | 2022-06-07 | Shanghai1 | Joy Zhou |  | BYTECH DONGTAI CO LTD | 146587 | 40 | VEN1230V1 | 4514851420 | BY012646634 |  |  | 100% with actual lot number on pack is BY012646634, should be BY204851420, the lot on carton is BY204851420. | Labeling | Accept by Deviation
67 | 2022-06-07 | Shanghai1 | Joy Zhou |  | BYTECH DONGTAI CO LTD | 146587 | 40 | VEN1235V1 | 4514851420 | BY012646634 |  |  | 100% with actual lot number on pack is BY012646634, should be BY204851420, the lot on carton is BY204851420. | Labeling | Accept by Deviation
68 | 2022-06-07 | Shanghai1 | Joy Zhou |  | BYTECH DONGTAI CO LTD | 146587 | 40 | VEN1235V1 | 4514948083 | BY205948083 |  |  | 100% with actual lot number on pack is BY012646634, should be BY205948083, the lot on carton is BY205948083. | Labeling | Accept by Deviation
69 | 2022-06-16 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | ISP-EU | NONELV200 | EU105966 | 22EBSP177 |  |  | There is AAMI level 2 on the neck label. | Visual | Accept by Deviation |  | Vendor corrected next production
70 | 2022-06-23 | Shanghai 2 | Peter Zhang |  | SHANGHAI M & G STATIONERY INC | 131715 | 22 | 136995 | 9000839529 | 71522050001 |  |  | All outer case with incorrect Part# and have one wrong word of item description,see as attached picture "136995 outer case"; | Visual | Rework |  | Training IQC and explain label reuqirments
71 | 2022-06-23 | Shanghai 2 | Peter Zhang |  | ZHEJIANG BEAUTY AND HEALTH ELECTRIC | 128805 | ISP-EU | CNB2 | 80522W22 | WZA033022B |  |  | wrong PO number on inner box label. And barcode number is inconsistent with lot number. | Labeling | Rework |  | Provide barcode training for vendor
72 | 2022-06-27 | Shanghai1 | Joy Zhou |  | WU XI LONG TENG TEXTILE CO LTD | 123001 | 50 | MDT211434PBP | 4514833753 | WTC22F |  |  | 100% Incorrect packaging:no polybag in box. | Visual | Rework
73 | 2022-06-28 | Shanghai 2 | Peter Zhang |  | SHANGHAI E WAY INDUSTRY CO LTD | 126422 | 82 | MSC282045 | 4514906256 | 42222050033 |  |  | The average weight out of tolerance. | Dimensional | Accept by Deviation |  | Re-explain PIC'S requirments for factory
74 | 2022-06-28 | Shanghai 2 | Joy Zhou |  | YANGZHOU PERFECT DAILY | 123460 | 81 | EQ21500 | 4600066258 | 50222060002 |  |  | 80w/ The item number on unit package(500ea) is EQ20500 instead of EQ21500 on spec. | Labeling | Accept by Deviation |  | Re-explain PIC'S requirments for factory
75 | 2022-06-29 | Shanghai 2 | Peter Zhang |  | AYK INTERNATIONAL INC | 155019 | 81 | MDTDBLTREADL | 4514993114 | AZD22050001 |  |  | ALL/Shipping carton labeling incorrect. ALL/The lot number on polybag is missing. | Labeling | Rework |  | Re-explain PIC'S requirments for factory
76 | 2022-07-01 | Pakistan | Muhammad Asad |  | LUCKY TEXTILES | 109756 | 50 | MDTSB8S30GRY | 4515038092 | R22E |  |  | Less weight | Dimensional | Rework |  | Replaced with a new one
77 | 2022-07-04 | Shanghai 2 | Joy Zhou |  | YANGZHOU PERFECT DAILY | 123460 | 33 | DYND30387 | 4515116092 | 55822070008 |  |  | Lot number on sterile package didn't match the lot on case. | Visual | Rework |  | Re-explain PIC'S requirments for factory
78 | 2022-07-07 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | ISP-EU | NONERP600L | XSP050622D | 22FBSP178 |  |  | Floding length is out of tolerance. SPEC: 38cm+/-2cm Defect: 42cm to 43cm. | Dimensional | Accept by Deviation |  | Vendor corrected next production
79 | 2022-07-12 | Shanghai1 | Joy Zhou |  | SUZHOU JINSHENG PLASTIC CO LTD | 113670 | ISP-Japan | GS-NC20FBR | SJP0526225 | SJP22070003J |  |  | Labeling Part# is incorrect. | Labeling | Accept by Deviation |  | Vendor used up old labeling and changed to new labeling.
80 | 2022-07-12 | Shanghai1 | Joy Zhou |  | SUZHOU JINSHENG PLASTIC CO LTD | 113670 | ISP-Japan | GS-NC40FBR | SJP0526225 | SJP22070004J |  |  | Labeling Part# is incorrect. | Labeling | Accept by Deviation |  | Vendor used up old labeling and changed to new labeling.
81 | 2022-07-12 | Shanghai1 | Joy Zhou |  | SUZHOU JINSHENG PLASTIC CO LTD | 113670 | ISP-Japan | NC10FBRJ | SJP0526225 | SJP22070002J |  |  | Labeling Part# is incorrect. | Labeling | Accept by Deviation |  | Vendor used up old labeling and changed to new labeling.
82 | 2022-07-12 | Shanghai1 | Joy Zhou |  | SUZHOU JINSHENG PLASTIC CO LTD | 113670 | ISP-Japan | NC30FSBRJ | SJP0526225 | SJP22070001J |  |  | Labeling Part# is incorrect. | Labeling | Accept by Deviation |  | Vendor used up old labeling and changed to new labeling.
83 | 2022-07-13 | NOC | Shawn Wei |  | ANDON HEALTH CO LTD | 129433 | 10 | MDS3002 | 4514993107 | 43322060003 |  |  | The actual cuff artwork is not the same with PIC. Old drawing was used | Labeling | Accept by Deviation |  | let vendor confirm with artwork team and align with their manufacturing
84 | 2022-07-15 | Hubei | Cherry Cheng |  | XIANTAO RAYXIN MEDICAL PRODUCTS | 116493 | 51 | CRI4001 | 4515141441 | 49322070249 |  |  | Folding are not accordance with PIC-00529 D-00017 | Visual | Accept by Deviation |  | Vendor corrected next production
85 | 2022-07-16 | Thailand | Chutima Ruensamran |  | COBES INDUSTRIES (MYANMAR) COMPANY LIMITED | 143305 | 21 | DYNJP8412 | 4515132062, 4515060431 | 30522070032 |  |  | Found the whole lot of products using incorrect pouch size (13.75” x 19.25”) The correct size is (20.75” x 18.25”) as per Medline BOM diagram | Visual | Rework
86 | 2022-07-23 | Pakistan | Muhammad Asad |  | LUCKY TEXTILES | 109756 | 50 | MDTR1011767D | 4515107882 | R22F |  |  | The size of the carton was not appropriate .The items after Packing were bulging out . | Visual | Rework |  | new carton with correct size
87 | 2022-07-26 | Shanghai 2 | Joy Zhou |  | YANGZHOU PERFECT DAILY | 123460 | 22 | 506653 | 9000839994 | 0332208A150 |  |  | 50w / The lot number Incorrect. 100% | Visual | Accept by Deviation |  | Re-explain PIC'S requirments for factory
88 | 2022-07-27 | NOC | Shawn Wei |  | HENAN YAOFENG MEDICAL SUPPLIES CO L | 153505 | 22 | 88601 | 9000840094 | 50522070003 |  |  | PO and lot number were printed wrong location. | Labeling | Rework |  | Check the case artwork and details carefully for each lot, QC pay more attention on raw material inspection.
89 | 2022-07-27 | NOC | Shawn Wei |  | HENAN YAOFENG MEDICAL SUPPLIES CO L | 153505 | 22 | 47285 | 9000840093 | 50522070003 |  |  | 100% PO and lot number were printed wrong location | Labeling | Rework |  | Check the case artwork and details carefully for each lot, QC pay more attention on raw material inspection.
90 | 2022-08-02 | Hefei | Echo Liu |  | LIANYUNGANG ANSON NON WOVEN PRODUC | 147612 | ISP-EU | SPTE-F8089 | KEN062922F | 22GAF3192 |  |  | 100% Arrow stamp excrescent on product without approved specification. | Visual | Rework |  | Update artwork
91 | 2022-08-05 | Hubei | Cherry Cheng |  | ALLMED INDUSTRIAL LIMITED | 22605 | 81 | NONEX0023F | 4515156073 | 6052208016 |  |  | The functional test don't complete, QA need NCMR to review. | Functional | Rework |  | PIC updated
92 | 2022-08-09 | NOC | Shawn Wei |  | LINGXIAN DONGSHENG TEXTILES CO LTD | 134330 | 71 | MSC9508LF | 4515173955 | 33022070001 |  |  | It should be 8roll per MC, but the labeling on Master Carton are printed as 12roll. Because the approved sticker vendor got is 12roll, so vendor wanted to ask for deviation but QA need them to rework before shipping. | Labeling | Rework |  | Training with vendor for more detail of our spec
93 | 2022-08-15 | TW | Carol Hsu |  | BESMED HEALTH BUSINESS CORP | 116085 | 29 | 121341 | 9000840476 | 08522080320 |  |  | Excessive flash on the surfaces interior | Visual | Rework |  | vendor 100% sort the lot
94 | 2022-08-18 | Hubei | Cherry Cheng |  | TOPWIDE HUBEI MEDICAL PRODUCTS | 137125 | 51 | NONCV150XXL | 4515226720 and 4515226723 | 12522071090 |  |  | Topwide produced with dimension 1’’ of storm flap , not in tolerance according to the latest diagram 1.5 +/-0.25 | Dimensional | Accept by Deviation |  | PIC updated
95 | 2022-08-18 | Hubei | Cherry Cheng |  | TOPWIDE HUBEI MEDICAL PRODUCTS | 137125 | 51 | NONCV150XXXXL | 4515226716 and 4515226722 | 12522071090 |  |  | Topwide produced with dimension 1’’ of storm flap , not in tolerance according to the latest diagram 1.5 +/-0.25 | Dimensional | Accept by Deviation |  | PIC updated
96 | 2022-08-20 | Hefei | Echo Liu |  | LIANYUNGANG ANSON NON WOVEN PRODUC | 147612 | ISP-EU | SPTE-NF8227 | KEN072822F | 22HAF3225 |  |  | Dimensions out of tolerance.The fenestration diameter should be 6cm. | Dimensional | Rework |  | Retrain worker
97 | 2022-08-24 | Hefei | Echo Liu |  | SINO PROTECTION HOLDINGS CO LTD | 105110 | 21 | DYNJP8304A | 4515320224 | 31122080029F |  |  | The size of breather pouch is a little different from Medline's BOM，rate 100% | Visual | Accept by Deviation |  | first sample confirmation and FQC
98 | 2022-08-24 | Shanghai 2 | Joy Zhou |  | SUZHOU DUKANGNING MEDICAL | 119653 | 22 | DYND30350 | 4515245550 | 0332208A150 |  |  | bag labeling missing lot. | Labeling | Rework |  | Re-explain QP-00019 requirments for factory
99 | 2022-08-25 | Hubei | Cherry Cheng |  | TOPWIDE HUBEI MEDICAL PRODUCTS | 137125 | 51 | NONCV150L | 4515226721/4515226718/4515226716 | 12522071090 |  |  | Storm flap width is not in tolerance according to the latest diagram 1.5 +/-0.25 | Dimensional | Accept by Deviation |  | PIC updated
100 | 2022-08-25 | Hubei | Cherry Cheng |  | TOPWIDE HUBEI MEDICAL PRODUCTS | 137125 | 51 | NONCV150XXXXL | 4515306440/4515306442 | 12522081260 |  |  | Storm flap width is not in tolerance according to the latest diagram 1.5 +/-0.25 | Dimensional | Accept by Deviation |  | PIC updated
101 | 2022-08-26 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | 51 | NONCV150XL | 4515306421 | 26322080825 |  |  | Storm flap width is 0.75'' , not in tolerance according to the latest diagram 1.5 +/-0.25. | Dimensional | Accept by Deviation |  | PIC updated
102 | 2022-08-26 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | ISP-EU | NONCV150XXXXL | 4515226717 | 26322080831 |  |  | Storm flap width is 1'' , not in tolerance according to the latest diagram 1.5 +/-0.25. | Dimensional | Accept by Deviation |  | PIC updated
103 | 2022-08-26 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | 51 | NONCV250XXL | 4515226717 | 26322080831 |  |  | Storm flap width is 2'' , not in tolerance according to the latest diagram 1.5 +/-0.25. | Dimensional | Accept by Deviation |  | PIC updated
104 | 2022-08-30 | TW | Carol Hsu |  | CAREMED SUPPLY INC | 126864 | 75 | MSC600907 | 4515226686 | 86422082001 |  |  | The hose length out of tolerance | Dimensional | Accept by Deviation |  | Vendor will update their internal spec. to reach our specification.
105 | 2022-08-30 | TW | Carol Hsu |  | CAREMED SUPPLY INC | 126864 | 75 | MSC600907 | 4515226686 | 86422082009 |  |  | The hose length out of tolerance,100% | Dimensional | Accept by Deviation |  | Vendor will update their internal spec. to reach our specification.
106 | 2022-09-02 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | 51 | NON25457M | 4515306436 | 26322080827 |  |  | Waist tie material is PP not as specification required that Ties will be same material as the gown. | Visual | Rework |  | Vendor corrected next production
107 | 2022-09-06 | Pakistan | Muhammad Asad |  | FEROZE1888 MILLS | 116221 | 50 | 3502BLKXS | 4515023413 | Z22C |  |  | Short Size of Back length | Dimensional | Rework |  | replacement of product
108 | 2022-09-07 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | 51 | NONCV150XL | 4515306425 | 26322080828 |  |  | Storm flap width is 1'' , not in tolerance according to the latest diagram 1.5 +/-0.25. | Dimensional | Accept by Deviation |  | PIC updated
109 | 2022-09-08 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | 51 | NONCV150XL | 4515306438 | 26322080830 |  |  | Storm flap width is 1'' , not in tolerance according to the latest diagram 1.5 +/-0.25. | Dimensional | Accept by Deviation |  | PIC updated
110 | 2022-09-08 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | 51 | NONCV250XXL | 4515306438 | 26322080830 |  |  | Storm flap width is 2'' , not in tolerance according to the latest diagram 1.5 +/-0.25. | Dimensional | Accept by Deviation |  | PIC updated
111 | 2022-09-10 | Pakistan | Muhammad Asad |  | MEED SKILLS SURGICAL | 124894 | 20 | 53791 | 110345046 | 894220908/1701-TAF |  |  | 100% Incorrect label | Labeling | Rework |  | Change the labels
112 | 2022-09-15 | Shanghai 2 | Joy Zhou |  | YANGZHOU PERFECT DAILY | 123460 | 50 | NON64WH | 4515293343 | 91122090001 |  |  | weight out of range | Dimensional | Accept by Deviation |  | Re-explain PIC'S requirments for factory
113 | 2022-09-15 | Hubei | Cherry Cheng |  | TOPWIDE HUBEI MEDICAL PRODUCTS | 137125 | 51 | NONCV150XL | 4515306446 | 12522081260 |  |  | Storm flap width is not in tolerance according to the latest diagram 1.5 +/-0.25 | Dimensional | Accept by Deviation |  | PIC updated
114 | 2022-09-16 | Shanghai1 | Joy Zhou |  | GCMEDICA ENTERPRISE LTD | 122183 | ISP-EU | TBGS-15-180 | GMD0628220 | 22HFG001 |  |  | 100% missing Net Weight and Gross Weight on label. | Labeling | Rework
115 | 2022-09-16 | Shanghai1 | Joy Zhou |  | GCMEDICA ENTERPRISE LTD | 122183 | ISP-EU | TBGS-15-300 | GMD0628220 | 22HFG001 |  |  | 100% missing Net Weight and Gross Weight on label. | Labeling | Rework
116 | 2022-09-16 | Shanghai1 | Joy Zhou |  | GCMEDICA ENTERPRISE LTD | 122183 | ISP-EU | TBGS-21-300-SE | GMD0628220 | 22HFG001 |  |  | 100% missing Net Weight and Gross Weight on label. | Labeling | Rework
117 | 2022-09-16 | Shanghai1 | Joy Zhou |  | GCMEDICA ENTERPRISE LTD | 122183 | ISP-EU | TBGS-21-360-SE | GMD0628220 | 22HFG001 |  |  | 100% missing Net Weight and Gross Weight on label. | Labeling | Rework
118 | 2022-09-16 | Shanghai1 | Joy Zhou |  | GCMEDICA ENTERPRISE LTD | 122183 | ISP-EU | TBGS-18-300 | GMD0628220 | 22HFG001 |  |  | 100% missing Net Weight and Gross Weight on label. | Labeling | Rework
119 | 2022-09-16 | Shanghai1 | Joy Zhou |  | GCMEDICA ENTERPRISE LTD | 122183 | ISP-EU | TBGS-18-360 | GMD0628220 | 22HFG001 |  |  | 100% missing Net Weight and Gross Weight on label. | Labeling | Rework
120 | 2022-09-22 | Hefei | Echo Liu |  | LIANYUNGANG ANSON NON WOVEN PRODUC | 147612 | ISP-EU | TB31301CE | KEN072822K | 22JDW251 |  |  | CSR Wrap incorrect size.Actual 60cm against spec. 76cm | Dimensional | Rework |  | Update artwork
121 | 2022-09-23 | Shanghai 2 | Joy Zhou |  | SUZHOU DUKANGNING MEDICAL | 119653 | 22 | 80957 | 9000841294 | 78322080001 |  |  | "Medline" was spelled wrongly on shipping marks. | Visual | Rework |  | Training IQC and explain label reuqirments
122 | 2022-09-23 | Shanghai 2 | Joy Zhou |  | SUZHOU DUKANGNING MEDICAL | 119653 | 22 | 94348 | 9000841294 | 78322080001 |  |  | "Medline" was spelled wrongly on shipping marks. | Visual | Rework |  | Training IQC and explain label reuqirments
123 | 2022-09-23 | Shanghai 2 | Joy Zhou |  | YANGZHOU PERFECT DAILY | 123460 | 18 | DYND14141MW | 4515195844 | 28722070003 |  |  | 100% ,125W/missing expiration date on PE bag | Visual | Accept by Deviation |  | Re-explain PIC'S requirments for factory
124 | 2022-09-23 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | 51 | NONCV250XXL | 4515393241 | 26322090835 |  |  | Storm flap width is 2'' , not in tolerance according to the latest diagram 1.5 +/-0.25. | Dimensional | Accept by Deviation |  | PIC updated
125 | 2022-09-23 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | 51 | NONCV250XL | 4515393241 | 26322090835 |  |  | Storm flap width is 2'' , not in tolerance according to the latest diagram 1.5 +/-0.25. | Dimensional | Accept by Deviation |  | PIC updated
126 | 2022-09-27 | Hubei | Cherry Cheng |  | XINTAO XINFA PLASTIC PRODUCTS LTD | 120542 | 51 | CRI1009 | 4515393249 | 54222091387 |  |  | 5 with Stitches per Inch (SPI) out of spec(21 24 29.20，21) | Dimensional | Accept by Deviation |  | Vendor corrected next production
127 | 2022-09-29 | Hefei | Echo Liu |  | LIANYUNGANG ANSON NON WOVEN PRODUC | 147612 | ISP-EU | TBAG3625-BULK | KEN082322L | 22JDF3264 |  |  | Drape incorrect size.Actual clear plastic panel size is 37.5inch different from spec 33.5inch, causing the final length over spec. | Dimensional | Rework |  | Update artwork
128 | 2022-09-29 | Shanghai 2 | Peter Zhang |  | AYK INTERNATIONAL INC | 155019 | 81 | MDTDBLTRXXXL | 4515416211 | AZQ22I00001 |  |  | ALL/The lot number on master carton and polybag is incorrect. | Visual | Rework |  | Training IQC and explain label reuqirments
129 | 2022-10-03 | Thailand | Chutima Ruensamran |  | COBES INDUSTRIES (MYANMAR) COMPANY LIMITED | 143305 | 21 | DYNJP9001 | 4515389026 ,4515389027 | 30522090025 |  |  | Incorrect-sized CSR wrap | Dimensional | Rework
130 | 2022-10-12 | Hefei | Echo Liu |  | LIANYUNGANG ANSON NON WOVEN PRODUC | 147612 | ISP-EU | TB29449A-BULK | KEN082322L | 22JJF3266 |  |  | Missing release paper cover | Visual | Rework |  | Retrain worker
131 | 2022-10-13 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | 51 | NONCV150XL | 4515393236 | 26322090836 |  |  | Storm flap width is 1'' , not in tolerance according to the latest diagram 1.5 +/-0.25. | Dimensional | Accept by Deviation |  | PIC updated
132 | 2022-10-13 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | 51 | NONCV250XL | 4515393240 | 26322090837 |  |  | Storm flap width is 2'' , not in tolerance according to the latest diagram 1.5 +/-0.25. | Dimensional | Accept by Deviation |  | PIC updated
133 | 2022-10-13 | Hubei | Cherry Cheng |  | HUBEI SAFEWAY PROTECTIVE | 113263 | 51 | NONCV250XXL | 4515393240 | 26322090837 |  |  | Storm flap width is 2'' , not in tolerance according to the latest diagram 1.5 +/-0.25. | Dimensional | Accept by Deviation |  | PIC updated
134 | 2022-10-18 | Shanghai1 | Joy Zhou |  | SUZHOU MECCAN IMP AND EXP CORP | 12911 | 18 | DYND16800 | 4515374910 | 91122100001 |  |  | 100% incorrect scanned UDI on inner boxes/bags labeling ,100%box labeling color not matched artwork provide by Medline | Labeling | Rework
135 | 2022-10-26 | Shanghai1 | Joy Zhou |  | JIANGSU CAINA MEDICAL CO LTD | 142897 | 10 | NON66160 | 4515370947 | 89722050001 |  |  | 100% incorrect barcode type on box. It should be GS1 Datamatrix,not Datamatrix | Labeling | Rework
136 | 2022-10-26 | Shanghai1 | Joy Zhou |  | JIANGSU CAINA MEDICAL CO LTD | 142897 | 10 | NON66160 | 4515416229 | 89722100001 |  |  | 100% incorrect barcode type on box. | Labeling | Rework
137 | 2022-10-27 | Shanghai 2 | Joy Zhou |  | YANGZHOU PERFECT DAILY | 123460 | 81 | MEDMAMCOMP1 | 4515416211 | AZQ22I00001 |  |  | Artwork of the inner box don’t match with actual product | Labeling | Accept by updating PIC |  | Re-explain PIC'S requirments for factory
138 | 2022-10-27 | Hubei | Cherry Cheng |  | XIANTAO DEMING HEALTHCARE PRODUCTS | 143706 | ISP-EU | CRIE4001B | XIA0826221 | 22KAD003 |  |  | Folding steps not according to approved specification | Visual | Accept by Deviation |  | QA confirmed that use the previous SPEC to do inspection
139 | 2022-10-28 | Shanghai 2 | Joy Zhou |  | SUZHOU DUKANGNING MEDICAL | 119653 | 15 | FLHC2 | SBL0916228 | SENE22W39 |  |  | Exp error | Labeling | Rework |  | Training IQC and explain label reuqirments
140 | 2022-10-31 | NOC | Shawn Wei |  | QINGDAO CHANGLONG | 120189 | ISP-Japan | SMDGVF05N-MIJ | QCS0915220 | QCS22105N44J |  |  | 100% case barcode were not readable | Labeling | Accept by Deviation |  | vendor will scan the barcode after the label was printed,no future lot was rejected again
141 | 2022-10-31 | NOC | Shawn Wei |  | QINGDAO CHANGLONG | 120189 | ISP-Japan | SMDGVF06N-MIJ | QCS0915220 | QCS22106N44J |  |  | 100% case barcode were not readable | Labeling | Accept by Deviation |  | vendor will scan the barcode after the label was printed,no future lot was rejected again
142 | 2022-11-08 | Shanghai 2 | Peter Zhang |  | ZHEJIANG RUNQIANG MEDICAL | 134969 | 22 | 131967 | 9000841835 | 96922090001 |  |  | All bags lacking seal line of fold after packed | Visual | Rework |  | Re-explain PIC'S requirments for factory
143 | 2022-11-08 | Shanghai 2 | Joy Zhou |  | SUZHOU DUKANGNING MEDICAL | 119653 | 22 | 131967 | 9000842202 | 96922100001 |  |  | all bags lacking seal line of fold after packed（as attached picture 01）. | Visual | Rework |  | Re-explain PIC'S requirments for factory
144 | 2022-11-16 | Shanghai 2 | Joy Zhou |  | SUZHOU DUKANGNING MEDICAL | 119653 | 60 | HEB150S | 4515476141 | 48822115001 |  |  | Expiration date is not required on the pack in PIC-00116 Revision 93 | Visual | Accept by updating PIC |  | Re-explain PIC'S requirments for factory
145 | 2022-11-16 | Shanghai1 | Joy Zhou |  | SUZHOU MECCAN IMP AND EXP CORP | 12911 | 20 | 58965 | 9000841931 | 91122110001 |  |  | height out of tolerance | Dimensional | Accept by Deviation |  | Update the drawing
146 | 2022-11-21 | Shanghai 2 | Peter Zhang |  | SHANGHAI KINDLY MEDICAL INSTRUMENT | 125341 | 22 | 145169 | 900084217 | 20221108 |  |  | Actual product is incorrect with the PIC. Dimension out of tolerance | Dimensional | Accept by updating PIC |  | Re-explain PIC'S requirments for factory
147 | 2022-11-23 | Shanghai 2 | Peter Zhang |  | ZHEJIANG WELLONG MEDICAL TECHNOLOGY | 148639 | 22 | 57404 | 9000842427 | 63922110800 |  |  | Paper width out of tolerance | Dimensional | Accept by Deviation |  | Re-explain PIC'S requirments for factory
148 | 2022-12-05 | Hubei | Cherry Cheng |  | ALLMED INDUSTRIAL LIMITED | 22605 | 20 | 56785 | 9000842485 | 6052212040 |  |  | Description in ETQ , this item is Non Index or Banded, it means there is no index and no band, but actually this item is banded. | Labeling | Accept by Deviation |  | ETQ description updated
149 | 2022-12-12 | Dongguan | Tim Li |  | XIAMEN ZHENGDONG ELECTRONICS CO LTD | 155800 | 10 | MDSSAFELT | 4702928513 | 202212MDL001 |  |  | 13W/Box size is wrong which it caused Its’ label is broken when put it on/100% | Labeling | Rework |  | Train QC to do actual fit test, not only check the size of all packaging materials.
150 | 2022-12-15 | NOC | Shawn Wei |  | INTCO MEDICAL HK CO LIMITED | 125123 | 40 | 6MSV601 | 4515627647 | ZIM212627647 |  |  | The color coded strip for S size should be orange, but vendor printed as Red. | Labeling | Rework |  | Training with vendor for more detail of our spec
151 | 2022-12-15 | Hubei | Cherry Cheng |  | XIANTAO DEMING HEALTHCARE PRODUCTS | 143706 | 51 | CRI4001 | 4515569121 | 70622120038 |  |  | All bags lacking seal line of fold after packed, 100% | Labeling | Rework |  | Vendor corrected next production
152 | 2022-12-16 | NOC | Shawn Wei |  | LINGXIAN DONGSHENG TEXTILES CO LTD | 134330 | 75 | MDT823330XL | 4515569030 | 33022110001 |  |  | weight out of spec. (higher than maximum of specification 680~690g) | Dimensional | Accept by Deviation |  | follow our latest spec when the stock runs out
153 | 2022-12-26 | Shanghai1 | Joy Zhou |  | GCMEDICA ENTERPRISE LTD | 122183 | ISP-EU | MDL1CE | GMD111022X | GMD23W02 |  |  | 100% wrong label on the carton.There should be no rectangle around the symbols. | Labeling | Rework
154 | 2022-12-27 | Shanghai 2 | Peter Zhang |  | ZHEJIANG WELLONG MEDICAL TECHNOLOGY | 148639 | 22 | 57404 | 9000842428 | 63922120800 |  |  | Paper width out of tolerance | Dimensional | Accept by Deviation |  | Re-explain PIC'S requirments for factory
155 | 2022-12-30 | NOC | Shawn Wei |  | LINGXIAN DONGSHENG TEXTILES CO LTD | 134330 | 34 | MDT828306M | 4515578486 | 33022120001 |  |  | Vendor didn't use updated labeling of RJ18LXT, after confirmed with QE, vendor need to rework with correct artwork. | Labeling | Rework |  | Training with vendor for more detail of our spec
156 | 2022-12-30 | Shanghai 2 | Joy Zhou |  | SUZHOU DUKANGNING MEDICAL | 119653 | 22 | 146926 146927 146929 146930 146931 | 9000843291 | 20221208 |  |  | Incorrect Product | Visual | Accept by updating PIC |  | N/A
157 | 2023-01-03 | Joy | Joy Zhou | Song Sun | CONOD MEDICAL CO LTD | 117969 | 18 | M102140B | 4600075329 | 96922120001 | 16 | 320 | 20W/missing courtesy tab on towels, 100% | visual | Deviation# 14573 | Vendor missed this requirement in PIC | Vendor update their internal spec
158 | 2023-01-03 | Joy | Joy Zhou | Song Sun | CONOD MEDICAL CO LTD | 117969 | 18 | M102140B | 4600075330 | 96922120001 | 16 | 320 | 20W/missing courtesy tab on towels, 100% | visual | Deviation# 14573 | Vendor missed this requirement in PIC | Vendor update their internal spec
159 | 2023-01-03 | Leo Gao | Leo Gao | Richard FU | COM BRIDGE INTERNATIONAL CO LTD | 119885 | 30 | TGLCANE | 4515554456 | 88523010001 | 250 | 3000 | The end of the inner pole doesn’t end at the middle of two adjacent height adjustment holes.100% | visual | Accept by Deviation | PIC is too strict. | Deviation #14521
160 | 2023-01-04 | Tran Cris | Cris Tran | Hoang Le | LOTUS CHEMICAL TECHNOLOGY CO LTD | 147176 | 81 | NON026370 | 4515700974 | 17623010003 | 108 | 54000 | 100% Different packaging configuration between PIC and actual product: PIC: 100ea/pk, 5pk/cs, total 500ea/cs Actual: 25ea/pk, 20pk/cs, total 500ea/cs | visual | QA accepted as is | Vendor missed to update their internal packing configuration (Medline already sent them PIC update information) | N/A, for future lots vendor must read the PIC carefully to avoid similar errors.
161 | 2023-01-10 | Peter | Peter Zhang | Bona | ZHEJIANG WELLONG MEDICAL TECHNOLOGY | 148639 | 22 | 507945 | 9000843334 | 63922121600 | 50 | 100000 | :have different packing quantities from ETQ. 400 Each/Case in ETQ; but actual is 2000 Each/Case, | visual | Accept by updating the PIC and ETQ | ETQ not updated timely
162 | 2023-01-10 | Peter | Peter Zhang | Bona | ZHEJIANG WELLONG MEDICAL TECHNOLOGY | 148639 | 22 | 507953 | 9000842427, 9000842428 | 63922121600 | 40 | 8000 | have different packing quantities from PIC-00286 & ETQ. 2000 Each/Case in PIC-00286 & ETQ; but actual is 200 Each/Case, [detail as 1ea/MSC2705 pouch ; 50 ea/MSC2705 box ; 200ea/cs(507053 label)] | visual | Accept by updating the PIC and ETQ | ETQ not updated timely
163 | 2023-01-10 | Peter | Peter Zhang | Bona | ZHEJIANG WELLONG MEDICAL TECHNOLOGY | 148639 | 22 | 507952 | 9000842427, 9000842428 | 63922121600 | 30 | 12000 | have different packing quantities from PIC-00286 & ETQ. 2000 Each/Case in PIC-00286 & ETQ,; but actual is 400 Each/Case, [detail as 1ea/MSC2703 pouch ; 100ea/MSC2703 box ; 400ea/cs (507952 label)] | visual | Accept by updating the PIC and ETQ | ETQ not updated timely
164 | 2023-01-10 | Peter | Peter Zhang | Bona | ZHEJIANG WELLONG MEDICAL TECHNOLOGY | 148639 | 22 | 507944 | 9000843334 | 63922121600 | 15 | 30000 | :have different packing quantities from ETQ. 200 Each/Case in ETQ; but actual is 2000 Each/Case, | visual | Accept by updating the PIC and ETQ | ETQ not updated timely
165 | 2023-01-11 | Joy | Joy Zhou | Ken Qin | ANHUI BYTECH MEDICAL SUPPLIES CO LT | 155283 | 40 | VEN1230PFV1 | 4515714937 | BY211521794 | 21 | 24192 | 32/w wrong LOT on polybag,100% | Labeling | Deviation | Misunderstanding between QA and vendor | Address and train Lot format with vendor again
166 | 2023-01-11 | Joy | Joy Zhou | Ken Qin | ANHUI BYTECH MEDICAL SUPPLIES CO LT | 155283 | 40 | VEN1235PFV1 | 4515714937 | BY211521794 | 17 | 19584 | 32/w wrong LOT number,100% | Labeling | Deviation | Misunderstanding between QA and vendor | Address and train Lot format with vendor again
167 | 2023-01-12 | Shawn Wei | Shawn Wei | Sam Wang | ABS INTERNATIONAL INC | 154286 | 50 | PM2127-18 | 4515604828 | 28623010001 | 670 | 8040ea | the dimensions have been updated from 26in (L) X 20in(W) to 19in X 24in, and vendor followed old specification in mass production. non-conformance rate 100% | Dimensional | Accept as is | the specification has been updated but vendor followed old specification to manufacture the products | Double check the updated specification and confirm with Medline representative if any confuse or misunderstanding about the update
168 | 2023-01-13 | Peter | Peter Zhang | Cristina | ZHEJIANG BEAUTY AND HEALTH ELECTRIC | 128805 | ISP-EU | CCB3 CCSB4 CNB2 CNB2 CUB1 DYND70880E DYND70885E | WZA110922R WZA110922R WZA110922R WZA112222E WZA112222E WZA110922R WZA1122221 | 80523W07 80523W07 80523W07 80523W08 80523W08 80523020002 | 3 3 7 4 45 72 2 | 3000 1500 7000 4000 45000 72000 2000 | All/ MFG.date and EXP.date printed on case and box label don’t reflect actual MFG.date and EXP.date. Printed:2023-02-01 2028-02-01 Should be: 2023-01-01 2028-01-01 | visual | Rework | IPQC not check the label information | Vendor corrected next production
169 | 2023-01-16 | Joy | Joy Zhou | Ken Qin | SUMEC TEXTILE COMPANY LIMITED | 153741 | 40 | MDTNC6S27 | 4515547062 | 74122L | 390 | 14040 | 13/w weight out of tolerance,100% | Dimensional | Deviation | Insufficient communication for new item criteria | Vendor follows PIC requirements for other shipments
170 | 2023-01-16 | Joy | Joy Zhou | Ken Qin | SUMEC TEXTILE COMPANY LIMITED | 153741 | 40 | MDTNC6S27 | 4515827779 | 74122L | 263 | 9468 | 13/w weight out of tolerance,100% | Dimensional | Deviation | Insufficient communication for new item criteria | Vendor follows PIC requirements for other shipments
171 | 2023-01-26 | Joy | Joy Zhou | Ken Qin | YANGZHOU PERFECT DAILY | 123460 | 22 | 500433 | 9000842903 | 46022120228 | 418 | 240768 | 8/w no expiration date on outer case, 100% | Labeling | Deviation | Incorrect artwork | Artwork updated with adding expiration date on outer case
172 | 2023-01-31 | Peter | Peter Zhang | Pennie | SHANGHAI KINDLY MEDICAL INSTRUMENT | 125341 | 55 | 146694 146695 | 9000843916 | 20230108 | 50 184 | 15000ea 55200ea | Po # on case label is 9000843428 instead of 9000843916 on spec (as attached picture 146694) Item 146695 Defect 1—8w / Po # on case label is 9000843428 instead of 9000843916 on spec (as attached picture 146695) Defect 2—13w / incorrect product description (as attached picture 146695 & 2) | Visual | Accept by updating the PIC and ETQ | ETQ not updated timely
173 | 2023-02-01 | Cherry | Cherry Cheng | Kelsey Li | WINNER MEDICAL CO LTD | 115279 | 20 | 508714 | 9000843141 | 27923010005 | 1600 | 1600000 | According to PIC-00550, item 508714 is bulk packaged without box. But vendor Winner produced this item with 20 boxes per case. | visual | Accepted by NCMR | Communication gap | Vendor comfirmed with Medline and changed package.
174 | 2023-02-07 | Cherry | Cherry Cheng | Kelsey Li | CROWN NAME DISPOSABLE HYGIENE PRODU | 147637 | 51 | NON27340 | 4515656711 | 63723010001 | 465 | 46500 | According to the latest diagram PIC-00528 D-00032, back of the gown is ultrasonic weld, but actually it is dotted line. Body seam is external weld, but actually it is inside weld. | Visual | Accepted by NCMR | Vendor not update in time | Vendor comfirmed with Medline and changed next order.
175 | 2023-02-07 | Peter | Peter Zhang | PeterZhang | NINGBO YICHENG COMMODITY CO LTD | 156003 | 81 | MDT217DISP1 | 323010001 | 4515779205 | 440 | 110000 | Width out of tolerance (actual 5.7”) All/Packing quantity incorrect Actual: 5bags/cs, 50ea/bag | Dimensional | Accept by updating the PIC and ETQ | ETQ not updated timely
176 | 2023-02-07 | Peter | Peter Zhang | Bona | HAINING HENGSHANG KNITTING CO LTD | 154006 | 50 | MDTIU8SEFBLU | 4515732903 | HX23A |  | 28224 | The LOT#HX23A is still used the old Lot Numbering System，it should be"LOT#00623A" whcih follow the new Lot Numbering System in PIC-00114. | visual | Accept by updating the PIC and ETQ | ETQ not updated timely
177 | 2023-02-09 | Cherry | Cherry Cheng | Judy Yang | WINNER MEDICAL CO LTD | 115279 | 21 | 508714 | 4600080709 | 27923020011 | 1600 | 1600000 | 100% rejection. According to PIC-00550, item 508714 should be bulk packaged without box. But vendor packed this item with 20 boxes per case. | visual | Accept by Deviation | Vendor ignore the package requirement in PIC. | Vendor corrected next production
178 | 2023-02-09 | Cherry | Cherry Cheng | Judy Yang | WINNER MEDICAL CO LTD | 115279 | 21 | 508714 | 4600080709 | 27923020011 | 1600 | 1600000 | 100% rejection. According to PIC-00550, item 508714 should be bulk packaged without box. But vendor packed this item with 20 boxes per case. | visual | Accept by Deviation | Vendor ignore the package requirement in PIC. | Vendor corrected next production
179 | 2023-02-16 | Cherry | Cherry Cheng | Kelsey Li | HUBEI SAFEWAY PROTECTIVE | 113263 | ISP-EU | CRIE2020 | XSP111122B | 22MASP190 | 192 | 38400 | 1.Case label is not the latest reversion.Vendor is LAB190727, latest is LAB225019. IFU is not the latest either. Vendor used IFU LAB180904 04/12/2018, the latest is IFU LAB180904 08/2022 2.8 shoe covers are founnd holes at the seam. | Labeling | Accepted by deviation | Vendor not update in time | Vendor will use latest label per the confirmation with EU
180 | 2023-02-16 | Cherry | Cherry Cheng | Kelsey Li | HUBEI SAFEWAY PROTECTIVE | 113263 | ISP-EU | CRIE2020 | XSP120922C | 23BASP193 | 500 | 100000 | Case label is not the latest reversion.Vendor is LAB190727, latest is LAB225019. IFU is not the latest either. Vendor used IFU LAB180904 04/12/2018, the latest is IFU LAB180904 08/2022 | Labeling | Accepted by deviation | Vendor not update in time | Vendor will use latest label per the confirmation with EU
181 | 2023-02-17 | Cherry | Cherry Cheng | Kelsey Li | HUBEI SAFEWAY PROTECTIVE | 113263 | ISP-EU | CRIE2102W | XSP120922A | 23B02SP193 | 13 | 13000 | Case label is not the latest reversion.Vendor is LAB190728-LAB190729, latest is CRIE2102W_LAB222682_LAB222681. IFU is not the latest either. Vendor used IFU LAB180904 04/12/2018, the latest is IFU LAB180904 08/2022. | Labeling | Accepted by deviation | Vendor not update in time | Vendor will use latest label per the confirmation with EU
182 | 2023-02-17 | Cherry | Cherry Cheng | Kelsey Li | HUBEI SAFEWAY PROTECTIVE | 113263 | ISP-EU | CRIE2102W | XSP120922C | 23B02SP194 | 63 | 63000 | Case label is not the latest reversion.Vendor is LAB190728-LAB190729, latest is CRIE2102W_LAB222682_LAB222681. IFU is not the latest either. Vendor used IFU LAB180904 04/12/2018, the latest is IFU LAB180904 08/2022. | Labeling | Accepted by deviation | Vendor not update in time | Vendor will use latest label per the confirmation with EU
183 | 2023-02-21 | Cherry | Cherry Cheng | Kelsey Li | WINNER MEDICAL CO LTD | 115279 | ISP-Japan | K3030-SF-10P-MIJ | WID1121221 | WID230100001J | 69 | 207000 | 1.Weight per piece is out of tolerance. SPEC:3.1±0.1g Defect:3.3g to 3.4 g 2.Folding dmension s are out of tolerance, please see below table for detail. | Dimensional | Accepted by deviation | Communication gap | Vendor will communicated with Japan and improve next time
184 | 2023-02-21 | Cherry | Cherry Cheng | Kelsey Li | WINNER MEDICAL CO LTD | 115279 | ISP-Japan | K3030-SF-MIJ | WID1121221 | WID230100001J | 12 | 36000 | 1.In PIC-00727 Produc tDescription, it is 10PK/CASE, 300E/PK,, but actually it is 20PK/CASE, 150EA/PK. 2.Weight Requirements (g/sqm) is out of tolerance. SPEC: 28.2 - 32.5 Defect: 32.8 3.Weight per piece is out of tolerance. SPEC:3.1±0.1g Defect:3.3g to 3.4 g 3.Folding dmension s are out of tolerance, please see below table for detail. | Dimensional | Accepted by deviation | Communication gap | Vendor will communicated with Japan and improve next time
185 | 2023-02-21 | Cherry | Cherry Cheng | Kelsey Li | WINNER MEDICAL CO LTD | 115279 | ISP-Japan | K3030-4F-MIJ | WID120922A | WID23020001J | 9 | 27000 | 1.Weight Requirements (g/sqm) is out of tolerance. SPEC: 28.2 - 32.5 Defect: 32.69 3.Weight per piece is out of tolerance. SPEC:3.1±0.1g Defect:3.3g to 3.4 g | Dimensional | Accepted by deviation | Communication gap | Vendor will communicated with Japan and improve next time
186 | 2023-02-21 | Peter | Peter Zhang | Garen | VR MEDICAL INC | 135063 | 10 | DYNJAAHMEHF | 6323010002 | 4515785841 | 143 | 7150 | All/ Incorrect barcode scanning | Visual | Rework | Vendor not scan the barcode before mass production
187 | 2023-03-01 | Joy | Joy Zhou | Ken Qin | SUMEC TEXTILE COMPANY LIMITED | 153741 | 50 | MDTPC6S34 | 4515516201 | SXU22K | 96 | 13824 | 8/w LOT format wrong,100% | Labeling | Deviation | Division changes LOT format without informing vendor | Vendor updates LOT format before deviation expired
188 | 2023-03-01 | Peter | Peter Zhang | allen | ZHEJIANG SHUYOU SURGICAL INSTRUMENT | 146095 | 10 | ESSHROUD1 ESSHROUD10 | 0000202355 | 09522060019 0000202355 |  |  | • The scanning format of the GS1 DataMatrix is incorrect. | Visual | Rework | Vendor not scan the barcode before mass production
189 | 2023-03-02 | Joy | Joy Zhou | Peter Lu | CITOTEST LABWARE MANUFACTURING CO L | 132415 | 33 | MLAB13SNPN | 4515900215 | 41523040001 | 2 | 40000 | 5w/ Diameter: Bottom Plug out of tolerance. 100% | Dimensional | rework | Use the wrong mold | double verify before production and QC need check first sample
190 | 2023-03-07 | Peter | Peter Zhang | Jeff | ANJI HONGDE MEDICAL PRODUCTS CO LTD | 148783 | 22 | 134347 134348 134349 134350 | 0000202357 | 20230214 | 1 3 6 82 59 | 576 864 1152 11808 5664 | All/ the color of all products are not consistent with the confirm sample, but the color of all products are consistent. | visual | Rework | sales department of Hongde not share the approved samples to their QC
191 | 2023-03-13 | Joy | Joy Zhou | Peter Lu | ZHANGJIAGANG WEIFENGRUN PLASTIC CO | 145419 | 40 | CLE2000HDPE | 4515850905 | WF315850905 | 1800 | 14400000 | Lot# on both carton & box should be WF303850905 instead of WF315850905. 100% | Labeling | Deviation | Supplier forget the lot format requirement after a long time no order | Have address the requirement to vendor and retrain
192 | 2023-03-13 | Tran Cris | Cris Tran | Hoang Le | LOTUS CHEMICAL TECHNOLOGY CO LTD | 147176 | 22 | 26900Y | 9000843189 | 17623010001 | 37 | 9250 | 100% Incorrect printing on actual product, product made from Vietnam vendor but printing said "made in China" | Labeling | QA accepted as is | Vendor received old diagram (make for China vendor) and didn't check it carefully, they just apply for mass production and didn't see it as error. | N/A, vendor has been sent a new artwork (Made in Vietnam) for future shipment.
193 | 2023-03-13 | Tran Cris | Cris Tran | Hoang Le | LOTUS CHEMICAL TECHNOLOGY CO LTD | 147176 | 22 | 509023Y | 9000843189 | 17623020002 | 51 | 12750 | 1. 100% Different packaging configuration between PIC and actual product PIC: 100ea/cs Actual: 250ea/cs (25ea/pk, 10pk/cs) 2. 100% Printing color different from artwork (actual products have lighter color) 3. 100% Width of product out of tolerance (spec 18 +/- 1’’), (actual 19.75’’) | Dimensional | Rework | Vendor received old diagram on PIC-00212 while on new project phase, after that Div QA updated these SKUs into new PIC-00587 (update packaging configuration, printing color and dimension) but vendor didn't check it, they just apply for mass production and didn't see it as error | Vendor scrap all the product and still not submit for re-inspection.
194 | 2023-03-15 | Joy | Joy Zhou | Song Sun | CONOD MEDICAL CO LTD | 117969 | 15 | DYNJWE1305 | 4515813268 | 96923030002 | 36 | 2160 | 15w/ crack marking on Silicone Reservoir cover ring, 75%; GTIN number is incorrect 100% | visual | rework | Injection molding issue, incoming inspection didn't detect this defect, Medline provided incorrrect artwork | Tighten incoming inspection. Medline provide updated artwork to vendor
195 | 2023-03-15 | Joy | Joy Zhou | Song Sun | CONOD MEDICAL CO LTD | 117969 | 15 | DYNJWE1305 | 4515813269 | 96923030002 | 18 | 1080 | 15w/ crack marking on Silicone Reservoir cover ring, 75%; GTIN number is incorrect 100% | visual | rework | Injection molding issue, incoming inspection didn't detect this defect, Medline provided incorrrect artwork | Tighten incoming inspection. Medline provide updated artwork to vendor
196 | 2023-03-15 | Joy | Joy Zhou | Song Sun | CONOD MEDICAL CO LTD | 117969 | 15 | DYNJWE1305 | 4515813270 | 96923030002 | 36 | 2160 | 15w/ crack marking on Silicone Reservoir cover ring, 75%; GTIN number is incorrect 100% | visual | rework | Injection molding issue, incoming inspection didn't detect this defect, Medline provided incorrrect artwork | Tighten incoming inspection. Medline provide updated artwork to vendor
197 | 2023-03-15 | Joy | Joy Zhou | Song Sun | CONOD MEDICAL CO LTD | 117969 | 15 | DYNJWE1305 | CSI010323P | 96923030002 | 54 | 3240 | 15w/ crack marking on Silicone Reservoir cover ring, 75%; GTIN number is incorrect 100% | visual | rework | Injection molding issue, incoming inspection didn't detect this defect, Medline provided incorrrect artwork | Tighten incoming inspection. Medline provide updated artwork to vendor
198 | 2023-03-15 | Joy | Joy Zhou | Song Sun | CONOD MEDICAL CO LTD | 117969 | 22 | 97783 | 9000843832 | 96923030002 | 20 | 1200 | 15w/ crack marking on Silicone Reservoir cover ring, 75%; GTIN number is incorrect 100% | visual | rework | Injection molding issue, incoming inspection didn't detect this defect, Medline provided incorrrect artwork | Tighten incoming inspection. Medline provide updated artwork to vendor
199 | 2023-03-15 | Joy | Joy Zhou | Song Sun | CONOD MEDICAL CO LTD | 117969 | 22 | 97783 | 9000843833 | 96923030002 | 13 | 780 | 15w/ crack marking on Silicone Reservoir cover ring, 75%; GTIN number is incorrect 100% | visual | rework | Injection molding issue, incoming inspection didn't detect this defect, Medline provided incorrrect artwork | Tighten incoming inspection. Medline provide updated artwork to vendor
200 | 2023-03-15 | Joy | Joy Zhou | Song Sun | CONOD MEDICAL CO LTD | 117969 | 22 | 97783 | 9000843834 | 96923030002 | 15 | 900 | 15w/ crack marking on Silicone Reservoir cover ring, 75%; GTIN number is incorrect 100% | visual | rework | Injection molding issue, incoming inspection didn't detect this defect, Medline provided incorrrect artwork | Tighten incoming inspection. Medline provide updated artwork to vendor
201 | 2023-03-16 | Joy | Joy Zhou | Ken Qin | CHANGZHOU MAJOR MEDICAL PRODUCTS | 140088 | 20 | 56452 | 9000843849 | 8823020005 | 95 | 410400 | 8/w improper packing 100% | Labeling | Deviation | Dual source has different packaging | Updated PIC to keep them consitent
202 | 2023-03-16 | Joy | Joy Zhou | Amy Lu | SUZHOU JINSHENG PLASTIC CO LTD | 113670 | 15 | DYNDCLO1001 | 4515831838 | 67023023825 | 250 | 2000 | 6w/dirt greater than 2 mm on the inside of the canister,30% ; 20w/Missing instruction sheet;100% | visual | rework | Insufficient communication for new item criteria | Pre-check before mass production
203 | 2023-03-16 | Joy | Joy Zhou | Amy Lu | SUZHOU JINSHENG PLASTIC CO LTD | 113670 | 15 | DYNDCLO1501 | 4515831838 | 67023023820 | 250 | 2000 | 4w/dirt greater than 2 mm on the inside of the canister, 20% 20w/Missing instruction sheet, 100% | visual | rework | Insufficient communication for new item criteria | Pre-check before mass production
204 | 2023-03-24 | Echo Liu | Echo Liu | Hunter Hou | LIANYUNGANG BRONSON NON WOVEN | 135947 | ISP-Japan | ZT-099TC | KEN011123B | 23BAW049 | 10 | 180 | This lot was rejected due to all samples defect with wrong location. | Labeling | Rework | folding cause the wrong location of stamp. | Change the folding procedure.
205 | 2023-03-24 | Leo Gao | Leo Gao | Panson Huang | COM BRIDGE INTERNATIONAL CO LTD | 119885 | 30 | G1-200MBUCKET | 4515909711 | 88523040001 |  | 50 | Missing 1 bucket lid. PIC-00244 state "1 Bucket and 1 bucket lid",100% | Labeling | Rework | Prodcut requirement was not completely confirmation. | Updated BOM and new product implement requirement.
206 | 2023-03-24 | Cherry | Cherry Cheng | Jody Zhang | TOPWIDE HUBEI MEDICAL PRODUCTS | 137125 | 51 | NONLV140 | 4515831773 | 12523020170 | 72 | 7200 | 100% Dimension "J" INTERNAL BACK ULTRASONIC WELD sealing is not as same as diagram | visual | Rework | New product for the vendor, misunderstand the design | Have updated in the re-made batch
207 | 2023-03-29 | Leo Gao | Leo Gao | Benjamin Pan | COM BRIDGE INTERNATIONAL CO LTD | 119885 | 30 | G1-200LX1 | 4515909673/4515909712 | 88523040001 | 510 | 510 | Lid cannot be installed in place. 100% | Functional | Rework | Lid deformation | Internal CAPA following
208 | 2023-03-29 | Cherry | Cherry Cheng | Judy Yang | XIANTAO DEMING HEALTHCARE PRODUCTS | 143706 | 51 | NONLV100XL | 4515831767 | 70623030003 | 822 | 82200 | 100% rejection. The actual seam seal type is heat seal. But per specs, the seam seal should be ultrasonic | visual | Accept by Deviation | Vendor ignore the seam requirement in PIC. | Vendor corrected next production
209 | 2023-03-29 | Cherry | Cherry Cheng | Judy Yang | XIANTAO DEMING HEALTHCARE PRODUCTS | 143706 | 51 | NONLV100XL | 4515831767 | 70623030003 | 822 | 82200 | 100% rejection. The actual seam seal type is heat seal. But per specs, the seam seal should be ultrasonic | visual | Accept by Deviation | Vendor ignore the seam requirement in PIC. | Vendor corrected next production
210 | 2023-03-30 | Joy | Joy Zhou | Kevin Qiu | JIANGSU NANFANG MEDICAL CO LTD | 122909 | 81 | EQU45246V1 | 4515954907 | 90923030001 | 60 | 28800 | 50w/Incorrect component, just find Knuckle bandages, and missing Fingertip bandages in box. 100% | visual | rework | Internal communication is not good, vendor's internal spec hasn't been updated. | Vendor update their internal spec
211 | 2023-03-30 | Leo Gao | Leo Gao | Smith Wang | COM BRIDGE INTERNATIONAL CO LTD | 119885 | 30 | MDS80318RW | 4515909712 | 88523040001 |  | 216 | The manual is printed on one side only, the other side is blank, missing the English instruction. 100% | visual | Rework | The manual vendor was missing printing one page. | Internal CAPA following
212 | 2023-03-30 | Peter | Peter Zhang | allen | ZHEJIANG SHUYOU SURGICAL INSTRUMENT | 146095 | ISP-Japan | ESRK3010-MIJ | ZHN0108232 | ZHN23030014J | 1 | 50 | 1. PO number was missing on Master carton label 2. 2 sterilization pouches with paper adhesion on the seal part after opening. Refer to Picture 1,2. | Visual | Rework | IPQC not check the label information beofore mass production | Provide training for vendor, and strengthen process checking
213 | 2023-03-30 | Peter | Peter Zhang | Bona | ZHEJIANG SHUYOU SURGICAL INSTRUMENT | 146095 | ISP-Japan | ES65-MIJ | ZHN23030016J ZHN23030015J ZHN23030019J ZHN23030030J | ZHN0108231 ZHN0108232 ZHN122722B ZHN0302230 | 2 1 6 2 | 400 200 1200 400 | PO number was missing on Master carton label.(as attached photo 001) | Visual | Rework | IPQC not check the label information beofore mass production | Provide training for vendor, and strengthen process checking
214 | 2023-03-30 | Peter | Peter Zhang | Bona | ZHEJIANG SHUYOU SURGICAL INSTRUMENT | 146095 | ISP-Japan | ESPB3010-MIJ ESRK3010-MIJ ESRK3010-MIJ | ZHN23030021J ZHN23030017J ZHN23030014J | ZHN122722B ZHN0108231 ZHN0108232 | 74 2 1 | 3700 100 50 | PO number was missing on Master carton label | Visual | Rework | IPQC not check the label information beofore mass production | Provide training for vendor, and strengthen process checking
215 | 2023-04-01 | Shawn Wei | Shawn Wei | Sam Wang | ABS INTERNATIONAL INC | 154286 | 50 | NON2439322 | 4515909708 | 28623030001 | 120 | 1440ea | vendor followed old specification in mass production non-conformance rate 100% | Dimensional | Accept as is | the specification has been updated but vendor followed old specification to manufacture the products | Double check the updated specification and confirm with Medline representative if any confuse or misunderstanding about the update
216 | 2023-04-03 | Joy | Joy Zhou | Peter Lu | ZHANGJIAGANG WEIFENGRUN PLASTIC CO | 145419 | 40 | CLE2000HDPE | 4515984381 | WF315984381 | 1800 | 14400000 | Incorrect Lot#. Found lot# on both carton & box should be WF303984381 instead of WF315984381. 100% | Labeling | Deviation | Supplier forget the lot format requirement after a long time no order | Have address the requirement to vendor and retrain
217 | 2023-04-07 | Cherry | Cherry Cheng | Kelsey Li | XIANTAO RAYXIN MEDICAL PRODUCTS | 116493 | ISP-EU | CRIE1004 | RAY011023A | 23BAR332 | 1168 | 584000 | 1.Dimension of stretched elastic length is larger than SPEC. SPEC:20" to 21" Defect:22'' to 23'' 2.Labeling of case and box is not the latest. labeling is the former revision CRIE1004_LAB190802_803, IFU LAB180904 04/12/2018. | Dimensional | Accepted by deviation | Vendor not update in time | Vendor poruced according to the latest document from next order
218 | 2023-04-07 | Peter | Peter Zhang | Cristina | ZHEJIANG BEAUTY AND HEALTH ELECTRIC | 128805 | ISP-EU | DYND70800UK | 4800007779 | 80523030001 | 6 | 160 | All/Incorrect GTIN number. Actual:0084389962940 Should be: 00884389962940 | Visual | Rework | Missing number "8", IPQC not check all the numbers one by one | Provide training for vendor, and strengthen process checking
219 | 2023-04-10 | Peter | Peter Zhang | allen | NATURECARE COSMETICS CO LTD | 145140 | 60 | MB1402 MB1404 | 4515941514 | 1402304NC2023062901 | 417 417 | 10008 10008 | actually no PO number on carton markings, refer to the attachment. | Visual | Accept by updating the PIC | ETQ not updated timely
220 | 2023-04-10 | Peter | Peter Zhang | summer | HANGZHOU LINAN SANXIN | 130650 | 82 | NATURESOFT913 | 4516000706 | 650230402 | 686 | 343000 | package weight out of spec Spec: ≥182g Actual: 134g, 135g, 131g, 140g, 132g, 135g, 132g, 140g, 135g, 133g, 136g, 134g, 136g. | Dimensional | Accept by updating the PIC | ETQ not updated timely
221 | 2023-04-12 | Echo Liu | Echo Liu | David Cui | SES HEALTHCARE PRODUCTS | 120311 | ISP-EU | 599CEB | SES020723D | 23CAEL113 | 12 | 1800 | 50 with wrong position of elastic band: The elastic band sewed on that sleeve is on the opposite side, not same side as described on the specification（the elastic band at each end is inside）. | visual | Rework | Didn't follow the spec. strictly | Review the whole procedure.
222 | 2023-04-19 | Leo Gao | Leo Gao | Leo Gao | COM BRIDGE INTERNATIONAL CO LTD | 119885 | 30 | MDS86420XWTL | 4516000724 | 88523050001 | 4 | 48 | The specification of the height rang is 37” ~ 46” (tolerance: +/-0.5”), but the lowest and highest heights of the product are about 29” and 38” respectively.,100% | Dimensional | Rework | The material collector takes the material without permission. | Improved the warehouse management requirement.
223 | 2023-04-19 | Tran Cris | Cris Tran | Hoang Le | DANAMECO MEDICAL JOINT STOCK | 152348 | 51 | NONCV980L | 4516024013 4516024015 | 34823040001 | 18 3 | 450 75 | 100% product used wrong tag, all product have "Isolation Gown AAMI Level 2" while they're Coverall product | Labeling | Rework | Vendor missed to check the details information on the labelling, their current process only checking the size of label and size of product on label | Vendor rework by removing the old label and replace it with correct one
224 | 2023-04-19 | Peter | Peter Zhang | Bona | SHANGHAI HUA YUE PACKAGING PRODS | 120680 | 30 | CUR200ALL | 4516037453 | 68023040002 | 48 | 5760 | insert with incorrect word “arm” ,the correct word should be “leg”(as attached picture 001) | visual | Rework | artwork incorrect | Provide training for vendor, and strengthen label checking
225 | 2023-04-19 | Peter | Peter Zhang | Anderson | SUZHOU SUNING UNDERPAD | 118842 | 82 | MSCE261220 MSCE261420 | SSA030223B | SSA23W16 | 840 210 | 84000 21000 | Overall Length out of spec Spec: 207.5-212.5cm Actual:216.2cm/217cm | Dimensional | Accept by QA | The setting parameters of Cutting machine are too large | Provide training for vendor, and strengthen process checking
226 | 2023-04-19 | Noorul | Noorul Hamsudin | Norzila | HARTALEGA SDN BHD | 116523 | 40 | 823-2 | 4515691931 | AN304691931 | 500 | 1446000 | 8 / 8 cases with H tape.100% | visual | Deviation | Internal document not updated according to PIC. | Request deviation from Division. Update internal document to reflect the changes in PIC.
227 | 2023-04-19 | Noorul | Noorul Hamsudin | Norzila | HARTALEGA SDN BHD | 116523 | 40 | 823-3 | 4515691931 | AN304691931 | 916 | 2688000 | 8 / 8 cases with H tape.100% | visual | Deviation | Internal document not updated according to PIC. | Request deviation from Division. Update internal document to reflect the changes in PIC.
228 | 2023-04-19 | Noorul | Noorul Hamsudin | Norzila | HARTALEGA SDN BHD | 116523 | 40 | 823-4 | 4515691931 | AN304691931 | 600 | 1746000 | 8 / 8 cases with H tape.100% | visual | Deviation | Internal document not updated according to PIC. | Request deviation from Division. Update internal document to reflect the changes in PIC.
229 | 2023-04-19 | Noorul | Noorul Hamsudin | Norzila | HARTALEGA SDN BHD | 116523 | 40 | 823-5 | 4515691931 | AN304691931 | 50 | 95000 | 8 / 8 cases with H tape.100% | visual | Deviation | Internal document not updated according to PIC. | Request deviation from Division. Update internal document to reflect the changes in PIC.
230 | 2023-04-27 | Peter | Peter Zhang | Jeff | ANJI HUIFENG SURGICAL DRESSINGS CO | 154318 | ISP-EU | 35154-EU 35155-EU | HUI0301231 HUI0301232 | HF230420 HF230420 | 2 4 | 840 960 | The size of width do not meet the PIC required PIC required: 35154-EU (5.5-6.5cm) 35155-EU (9.0-11.0cm) Actual found: For 35154-EU: (4.8cm/5.0cm/5.3cm/4.7cm/5.0cm) For 35155-EU(7.0cm/7.2cm/7.0cm/6.8cm/7.2cm) | Dimensional | Accept by QA | The setting parameters of Cutting machine are too large | Provide training for vendor, and strengthen process checking
231 | 2023-04-28 | Dina | K. Dinakaran | Jayaprakash | TEXBOND NONWOVENS | 152093 | 51 | CRI1004 | 4515918646 | 9323040005 | 118 | 59000 | pieces Shortage-100% | visual | Rework | manual counting error | Manual counting error , awareness given to packing workers and ask them to add 2 pieces extra during packing to avoid shortage in future.
232 | 2023-05-08 | Leo Gao | Leo Gao | Benjamin Pan | A AND J HOLDING CO LTD | 112832 | 30 | MDS88200DML | 4516000627 | 83223050001 | 2 | 80 | the packaging didn't match PIC' spec,the supplier packs 10ea per case, but PIC requires 50ea per case,100% | Labeling | Rework | The artwork of shipper carton was wrong. | Updated the artwork.
233 | 2023-05-08 | Peter | Peter Zhang | allen | HANGZHOU LINAN SANXIN | 130650 | 82 | ULTRASOFT1013 | 650230502 650230503 | 4516086086 | 1120 200 | 560000 100000 | Product length is out of tolerance. Specs: length 30±1.5cm. Actual found 32.5, 32.1, 32.5, 32.3, 32.5, 32.8, 33, 32.4, 32.6, 32.8, 32.5, 32.6, 33.2 | Dimensional | Accept by Deviation | The setting parameters of Cutting machine are too large | Provide training for vendor, and strengthen process checking
234 | 2023-05-09 | Chutima | Chutima Ruensamran | May Lwin (mlwin) | COBES INDUSTRIES (MYANMAR) COMPANY LIMITED | 143305 | 51 | NONLV200XXL | 4516100659 | 30523050048 | 4 | 400 | The neck closure tape tab is placed at the right back panel of the gown which is different from PIC-00528 specification in which the tape tab should be at the left back panel. | visual | NCMR-0010259,deviation,Visual | Medline product manger(Maggie Wang) asked to put the tab tap at the right through email, but when we inspect, the spec does not changed yet, so when we asked the Div QA , they thought remainning the tape tab at the left is better for some reason and decide not to change the spec ,gave deviation for those effective lot | Vendor will follow the previous practice which it is refer to the product specification sheet PIC-00528 from the next production.
235 | 2023-05-09 | Chutima | Chutima Ruensamran | May Lwin (mlwin) | COBES INDUSTRIES (MYANMAR) COMPANY LIMITED | 143305 | 51 | NONLV200 | 4516100659 | 30523050036 | 400 | 40000 | The neck closure tape tab is placed at the right back panel of the gown which is different from PIC-00528 specification in which the tape tab should be at the left back panel. | visual | NCMR-0010269,deviation,Visual
236 | 2023-05-11 | Chutima | Chutima Ruensamran | May Lwin (mlwin) | COBES INDUSTRIES (MYANMAR) COMPANY LIMITED | 143305 | 51 | NONLV200XL | 4516100659 | 30523050030 | 560 | 56000 | The neck closure tape tab is placed at the right back panel of the gown which is different from PIC-00528 specification in which the tape tab should be at the left back panel. | visual | NCMR-0010287,deviation,Visual
237 | 2023-05-11 | Chutima | Chutima Ruensamran | May Lwin (mlwin) | COBES INDUSTRIES (MYANMAR) COMPANY LIMITED | 143305 | 51 | NONLV200 | 4516100659 | 30523050036 | 600 | 60000 | The neck closure tape tab is placed at the right back panel of the gown which is different from PIC-00528 specification in which the tape tab should be at the left back panel. | visual | NCMR-0010288,deviation,Visual
238 | 2023-05-12 | Asad | Muhammad Asad | Murad Younas | MOUNT BLONCK TRADING | 127094 | 20 | 5030 | 110343497 | 110343497-5030 | 15 | 15000 | Different lot number was pasted on cartons, 100% | Labeling | Rework | wrong lot # | changed the carton stikers
239 | 2023-05-12 | Leo Gao | Leo Gao | Panson Huang | BLISS HEALTH PRODUCTS CO LTD | 148948 | 30 | G3-600KRX1 | 4515941509 | 94823050001 | 1370 | 1370 | There is not lot # on retail box,100% | Labeling | Accept As is | New product communication wrong. | Accept As is
240 | 2023-05-12 | Leo Gao | Leo Gao | Panson Huang | BLISS HEALTH PRODUCTS CO LTD | 148948 | 30 | G2-611KRX1 | 4515941509 | 94823050002/94823050001 | 1370 | 1370 | There is not lot # on retail box,100% | Labeling | Accept As is | New product communication wrong. | Accept As is
241 | 2023-05-15 | Chutima | Chutima Ruensamran | KoKo Lin (klin) | COBES INDUSTRIES (MYANMAR) COMPANY LIMITED | 143305 | 51 | NONLV200 | 4516204843 | 30523050165 | 98 | 9800 | The neck closure tape tab is placed at the right panel of the gown which is different from PIC-000528 specification in which the tape tab should be at the left back panel. | visual | NCMR-0010301,deviation,Visual
242 | 2023-05-16 | Chutima | Chutima Ruensamran | Htaik Thu Aung (haung) | COBES INDUSTRIES (MYANMAR) COMPANY LIMITED | 143305 | 51 | NONLV200 | 4516100663 | 30523050080 | 300 | 30000 | The neck closure tape tab is placed at the right back panel of the gown which is different from PIC-00528 specification in which the tape tab should be at the left back panel. | visual | NCMR-0010311,deviation,Visual
243 | 2023-05-22 | Joy | Joy Zhou | Ken Qin | NANJING 3H MEDICAL PRODUCTS | 125843 | 20 | 56453 | 9000845476 | 84323050404-A | 14 | 60480 | 8 W/ Improper Packing, 100% | visual | Deviation | Division changes packing method without informing vendor in advance | Vendor updates packing method based on PIC requirements
244 | 2023-05-24 | Leo Gao | Leo Gao | Damon Mu | ZHENJIANG ASSURE MEDICAL EQUIPMENT | 148164 | 30 | K4186N44E1 | 4516128329 | 16423050001 | 24 | 24 | Seat extender tube OD is larger than seat tube ID.That caused the extender can't be installed. 100% | Functional | Rework | Worker used wrong tube. | Replace all extender tube with right size.
245 | 2023-05-24 | Leo Gao | Leo Gao | Benjamin Pan | ZHONGSHAN BLISS MEDICAL | 124956 | 30 | MDS80618T | 4516085980 | 95623050003 | 54 | 54 | the manual date didn't match PIC' SPEC. Product sheet show the date was 21/02/23, PIC'S spec showed 24/12/19 ,100% | Labeling | QA Accept As Is | The PIC didn't updated in time. | Updated the PIC
246 | 2023-05-24 | Tran Cris | Cris Tran | Binh Tran | LOTUS CHEMICAL TECHNOLOGY CO LTD | 147176 | 22 | 48202Y | 9000845197 | 17623040001 |  |  | 100% PO and lot number were printed wrong location | Labeling | Rework |  | The person printing the information checks carefully before printing, and the person packaging checks again before sealing the corrected location PO# and Lot# on cartons.
247 | 2023-05-31 | Tran Cris | Cris Tran | Hoang Le | DIEU THUONG JOINT STOCK COMPANY | 151367 | 81 | NON24272 | 4516100749 | 36723040003 | 80 | 80000 | 100% Carton have incorrect PO#, correct PO only have 10 digits, actual found with 11 digits | Labeling | Rework | IQC didn’t check the information on the label clearly, lead to findings FQC didn't spot out the error due to it's the same person with IQC | Vendor rework by putting correct sticker into cartons.
248 | 2023-06-01 | Joy | Joy Zhou | JANET XIA | SUZHOU JINGLE POLYMER MEDICAL | 118486 | 35 | 399G0025 | HD23051101 | 48623050001 | 16 | 16000 | 125w/incorrect centimeter marked on catheter .It should be 6-26cm instead of 6-28cm,100% | visual | rework | PIC hasn't been updated | Updated PIC
249 | 2023-06-02 | Shawn Wei | Shawn Wei | Sam Wang | ABS INTERNATIONAL INC | 154286 | 50 | NON24390 | 4516091802 | 28623E | 90 | 1080ea | the hemming thread color should be WHITE, but blue thread was used on the products. non-conformance rate 100% | visual | Rework | The sewing worker didn’t change to white thread for NON24390 after finishing the sewing of PM1824-15. And the inspector didn’t find out when inline/final inspection | 1. Strengthen quality management and keep training with their sewing workers and inspectors. 2. Taking cross inspections for all the products. 3. Daily checking all the materials before production start.
250 | 2023-06-05 | Tran Cris | Cris Tran | Hoang Le | LOTUS CHEMICAL TECHNOLOGY CO LTD | 147176 | 22 | 92047Y | 9000845198 | 17623040001 | 94 | 23500 | 100% Current artwork for item isn't correct (not up to dated) Old artwork: RK18LOT, New artwork: RG22LOT | Labeling | QA accepted as is | Vendor didn’t receive artwork change notification for 92047Y (Div22) even when it’s the same item with NON026330 (Div81), so they didn’t make changes | N/A, for future lots vendor will use the same artwork with Div81 product (the most updated one).
251 | 2023-06-07 | TW | Carol Hsu | Becky Song (bsong) | FORTECHGRP CO LTD | 150871 | 10 | MDS8701 | 4516176786 | 87123050001 | 109 | 545000 | 100% Package with print stains | Visual | Rework | stain by ink of the box | 100% sort the lot and change box, in the future vendor will apply for controling change to replacing the box.
252 | 2023-06-07 | Tran Cris | Cris Tran | Binh Tran | THANH DUNG TRANSPORT SERVICE ONE | 153715 | 82 | NATURESOFT913 | 202345 | 71523030001 |  |  | 1/The top and bottom edges of carton are deformed.100% 2/ Lot# are print incorrect position as requirement. 100% | Labeling | Rework |  | will update rework (not reinspection yet)
253 | 2023-06-08 | Dina | K. Dinakaran | Mohamed Maideen | SHINRAI TEXTILE PRIVATE LIMITED | 155832 | 51 | NONCV120XXL | 4516100716 | 83223050008 | 15 | 375 | 100% Inspected found Bad odour that smells Smokey. | Visual | Rework | It is suspected that, smells from fabric was there while fabric receive. Since all workers were wearing the mask for safety purposes, no one from vendor side were smelled the Smookey odor. | Fabric Manufactures were alerted for the smookey fabric issues to take the corrective actions.Shinrai will smell the fabric upon receiving for odor check.Also while Fabric Inspection. | The identified product were removed from Shipment. Will go in a following shipment
254 | 2023-06-08 | Dina | K. Dinakaran | Mohamed Maideen | SHINRAI TEXTILE PRIVATE LIMITED | 155832 | 51 | NONCV1205XL | 4516100716 | 83223050009 | 5 | 125 | 100% Inspected found Bad odour that smells Smokey. | Visual | Rework | It is suspected that, smells from fabric was there while fabric receive. Since all workers were wearing the mask for safety purposes, no one from vendor side were smelled the Smookey odor. | Fabric Manufactures were alerted for the smookey fabric issues to take the corrective actions.Shinrai will smell the fabric upon receiving for odor check while Fabric Inspection. | The identified product were removed from Shipment. Will go in a following shipment
255 | 2023-06-08 | Cherry | Cherry Cheng | Kelsey Li | XIANTAO RAYXIN MEDICAL PRODUCTS | 116493 | ISP-EU | FS65450 | RAY050423H | 23DAR387 | 560 | 672000 | Lot# is not correct, manufacture date 05/2023, lot# is 23DAR387. It should be 23Ear387. | Labeling | Accepted by deviation | Production defect | Vendor corrected the next order.
256 | 2023-06-09 | Chutima | Chutima Ruensamran | KoKo Lin (klin) | SINO PROTECTION YANGON NON WOVEN | 150832 | 22 | SPT-2206SMSMY | 9000845258 | 83223050006 | 6 | 600 | Part of the stamp (X-Long) is missing for the whole lot | Labeling | NCMR-0010471,rework,Labeling | 1. The inspector QA and QC are both new employees. Due to the inspector's negligence and lack of training in place, the stamping issue were not found and serious mistakes were made. 2. Team leader did not do the first inspection . | 1.Retrain stamping workers and relate staff follow the procedure , approved sample review , 1st piece inspection. 2, For quality and management personnel to conduct the correct training of production process control , inspection each process must be accurate .
257 | 2023-06-09 | Chutima | Chutima Ruensamran | KoKo Lin (klin) | SINO PROTECTION YANGON NON WOVEN | 150832 | 22 | SPT-2206SMSMY | 9000845304 | 83223050006 | 2 | 200 | Part of the stamp (X-Long) is missing for the whole lot | Labeling | NCMR-0010471,rework,Labeling
258 | 2023-06-09 | Cherry | Cherry Cheng | Kelsey Li | CROWN NAME DISPOSABLE HYGIENE PRODU | 147637 | ISP-EU | CRIE2010 | COW051023A | 23FBZO001 | 98 | 98000 | Barcode is not correct. | Labeling | Rework | Production defect
259 | 2023-06-09 | Cherry | Cherry Cheng | Kelsey Li | CROWN NAME DISPOSABLE HYGIENE PRODU | 147637 | ISP-EU | CRIE2010 | COW050423B | 23FBZO002 | 588 | 588000 | Barcode is not correct. | Labeling | Rework | Production defect
260 | 2023-06-09 | Cherry | Cherry Cheng | Kelsey Li | CROWN NAME DISPOSABLE HYGIENE PRODU | 147637 | ISP-EU | NONE35LKIT | COW050423B | 23FAZO002 | 540 | 27000 | According to SPEC of PL9110P, there is a pocket on the pant, actually there is no pocket. | visual | Accepted by deviation | Vendor not update in time | Vendor will use latest label per the confirmation with EU
261 | 2023-06-09 | Cherry | Cherry Cheng | Kelsey Li | CROWN NAME DISPOSABLE HYGIENE PRODU | 147637 | ISP-EU | NONE352XLKIT | COW050423B | 23FAZO002 | 24 | 1200 | According to SPEC of P2XL9110P, there is a pocket on the pant, actually there is no pocket. | visual | Accepted by deviation | Vendor not update in time | Vendor will use latest label per the confirmation with EU
262 | 2023-06-13 | Leo Gao | Leo Gao | Damon Mu | ZHENJIANG ASSURE MEDICAL EQUIPMENT | 148164 | 30 | MDS808200BNP | 4515789220 | 16423040001 | 102 | 102 | The weight capacity in spc is 250lbs, but in manual the weight capicity is 300lbs. 100% | visual | Rework | The artwork of Manual was wrong. | Updated manual artwork.
263 | 2023-06-15 | Leo Gao | Leo Gao | Kevin Hou | ZHENJIANG ASSURE MEDICAL EQUIPMENT | 148164 | 30 | MDSG4BSK | 4516132025 | 16423060001 | 2 | 40 | 100% Barcode format error,100% | Labeling | Rework | Tier 2 supplier used the barcode in the wrong format. | check the label before production
264 | 2023-06-16 | Shawn Wei | Shawn Wei | Sam Wang | JIANGYIN HANRONG TECHNOLOGY CO LTD | 123283 | 50 | MDTST4A27MIS | 4516086106 | 28323D | 3 | 900ea | vendor manufactured the products according to old specification, non-conformance rate 100% | Dimensional | Accept as is | the specification has been updated but vendor followed old specification to manufacture the products | Double check the updated specification and confirm with Medline representative if any confuse or misunderstanding about the update
265 | 2023-06-19 | Shawn Wei | Shawn Wei | Sam Wang | JIANGYIN HANRONG TECHNOLOGY CO LTD | 123283 | 50 | MDTST4A27MIS | 4516131908 | 28323E | 5 | 1500ea | vendor manufactured the products according to old specification, non-conformance rate 100% | Dimensional | Accept as is | the specification has been updated but vendor followed old specification to manufacture the products(stock) | Double check the updated specification and confirm with Medline representative if any confuse or misunderstanding about the update
266 | 2023-06-19 | Cherry | Cherry Cheng | Judy Yang | CROWN NAME DISPOSABLE HYGIENE PRODU | 147637 | 51 | NONSW1765S | 4516204850 | 63723060008 | 2 | 60 | 100% rejection. the first poly bag were not sealed and outer poly liner were not taped or tied. But per specs, frocks will be packaged in 2 inner poly liner/case liners. The first poly bag must be sealed, while the outer poly liner can be taped or tied | visual | Accept by Deviation | Vendor ignore the package requirement in PIC. | Vendor rework the package
267 | 2023-06-19 | Cherry | Cherry Cheng | Judy Yang | CROWN NAME DISPOSABLE HYGIENE PRODU | 147637 | 51 | NONSW1765L | 4516204851 | 63723060009 | 126 | 3780 | 100% rejection. the first poly bag were not sealed and outer poly liner were not taped or tied. But per specs, frocks will be packaged in 2 inner poly liner/case liners. The first poly bag must be sealed, while the outer poly liner can be taped or tied | visual | Accept by Deviation | Vendor ignore the package requirement in PIC. | Vendor rework the package
268 | 2023-06-19 | Cherry | Cherry Cheng | Judy Yang | CROWN NAME DISPOSABLE HYGIENE PRODU | 147637 | 51 | NONSW1765S | 4516204851 | 63723060009 | 1 | 30 | 100% rejection. the first poly bag were not sealed and outer poly liner were not taped or tied. But per specs, frocks will be packaged in 2 inner poly liner/case liners. The first poly bag must be sealed, while the outer poly liner can be taped or tied | visual | Accept by Deviation | Vendor ignore the package requirement in PIC. | Vendor rework the package
269 | 2023-06-19 | Cherry | Cherry Cheng | Judy Yang | CROWN NAME DISPOSABLE HYGIENE PRODU | 147637 | 51 | NONSW1765S | 4516204850 | 63723060008 | 2 | 60 | 100% rejection. the first poly bag were not sealed and outer poly liner were not taped or tied. But per specs, frocks will be packaged in 2 inner poly liner/case liners. The first poly bag must be sealed, while the outer poly liner can be taped or tied | visual | Accept by Deviation | Vendor ignore the package requirement in PIC. | Vendor rework the package
270 | 2023-06-19 | Cherry | Cherry Cheng | Judy Yang | CROWN NAME DISPOSABLE HYGIENE PRODU | 147637 | 51 | NONSW1765L | 4516204851 | 63723060009 | 126 | 3780 | 100% rejection. the first poly bag were not sealed and outer poly liner were not taped or tied. But per specs, frocks will be packaged in 2 inner poly liner/case liners. The first poly bag must be sealed, while the outer poly liner can be taped or tied | visual | Accept by Deviation | Vendor ignore the package requirement in PIC. | Vendor rework the package
271 | 2023-06-19 | Cherry | Cherry Cheng | Judy Yang | CROWN NAME DISPOSABLE HYGIENE PRODU | 147637 | 51 | NONSW1765S | 4516204851 | 63723060009 | 1 | 30 | 100% rejection. the first poly bag were not sealed and outer poly liner were not taped or tied. But per specs, frocks will be packaged in 2 inner poly liner/case liners. The first poly bag must be sealed, while the outer poly liner can be taped or tied | visual | Accept by Deviation | Vendor ignore the package requirement in PIC. | Vendor rework the package
272 | 2023-06-20 | Cherry | Cherry Cheng | Judy Yang | XIANTAO DEMING HEALTHCARE PRODUCTS | 143706 | ISP-EU | CRIE4001B | DMG050423A | 23FAD006 | 1338 | 66900 | 100% rejection. Missing PO number, gross weight, net weight on case labeling | Labeling | Rework | Vendor ignore the labeling requirement in PIC. | Vendor rework the labeling
273 | 2023-06-20 | Cherry | Cherry Cheng | Judy Yang | XIANTAO DEMING HEALTHCARE PRODUCTS | 143706 | ISP-EU | CRIE4001B | DMG050423A | 23FAD006 | 1338 | 66900 | 100% rejection. Missing PO number, gross weight, net weight on case labeling | Labeling | Rework | Vendor ignore the labeling requirement in PIC. | Vendor rework the labeling
274 | 2023-06-23 | Tran Cris | Cris Tran | Melody Phan | DIEU THUONG JOINT STOCK COMPANY | 151367 | 15 | DYNJE63654S | 4516210334 | 36723050003 | 165 | 4125 | Observed 100% pouches with incorrect REF "DYNJE63654R", correct should be DYNJE63654S. | Labeling | Rework | The issue detected was only 1 character between R and S, some SKUs have R in the end and some have S, the supplier made a mistake printing and QC team also did not detect to remove defects. | Vendor paste the sticker with the correct SKU on the wrong SKU position The firm re-train for related employees about Labeling Control Procedure to handle and inspect labels properly.
275 | 2023-06-26 | Peter | Peter Zhang | Allen | ZHEJIANG SHUYOU SURGICAL INSTRUMENT | 146095 | 15 | ESE15526 | 4516227373 | 9523060016 | 2 | 800 | Black PTFE was required in diagram in QUAL-96006, rev 3. But actually no black PTFE on product. Refer to attachment-‘product didn’t match the diagram’. | visual | Accept by updating PIC | ETQ not updated timely
276 | 2023-06-26 | Peter | Peter Zhang | Cristina | ZHEJIANG BEAUTY AND HEALTH ELECTRIC | 128805 | ISP-ANZ | DYND70802AU | WZA042623F | 80523060001 | 10 | 200 | Wrong certification number on the plug. See attached photo01 PIC spec: E1310 Actual: E3211 | visual | Accept by updating PIC | ETQ not updated timely
277 | 2023-06-30 | Echo Liu | Echo Liu | Lovick Li | SINO PROTECTION HOLDINGS CO LTD | 105110 | 21 | DYNJP4131 | 4516210359 | 11023070068 | 135 | 1620 | 5pcs with dimension out of tolerance. The length and width of the drape: our drawing are 125” and 91”,45”,23”. The actual products are 145”and 83”,40”,21.5”. | Dimensional | Rework | Vendor didn't produce per the latest SPEC. | Check the latest spec.
278 | 2023-07-03 | Noorul | Noorul Hamsudin | Halimatun | HARTALEGA SDN BHD | 116523 | 40 | MDS8085 | 4600086220 | AN307086220 | 300 | 360000 | 8 / 8 cases with wrong glove quantity on artwork (1,200 instead of 1,500).100% defect rate | Labeling | Rework | Overlook by QA Incoming during inspection | Re-order new packaging material.
279 | 2023-07-03 | Noorul | Noorul Hamsudin | Halimatun | HARTALEGA SDN BHD | 116523 | 40 | MDS8086 | 4600086220 | AN307086220 | 50 | 60,00 | 8 / 8 cases with wrong glove quantity on artwork (1,200 instead of 1,500).100% defect rate | Labeling | Rework | Overlook by QA Incoming during inspection | Re-order new packaging material.
280 | 2023-07-06 | Leo Gao | Leo Gao | Smith Wang | COM BRIDGE INTERNATIONAL CO LTD | 119885 | 30 | CVS980711V1 | 4516221916 | 88523070002 |  | 168 | The description for grip on shipper carton labeling is inconsistent with PIC, actual product and manual. 100% | Labeling | QA Accept As Is | The artwork of shipper carton was wrong. | QA Accept As It, update the artwork
281 | 2023-07-07 | Cherry | Cherry Cheng | Jody Zhang | HUBEI SAFEWAY PROTECTIVE | 113263 | ISP-EU | FS63150A | XSP05303EA | 23FASP205 | 75 | 45000 | 100% boxes were printed “Made in (country)”, which should be “Made in China”. | Labeling | Rework | The vendor forgot to update the area | The vendor will double check in next run
282 | 2023-07-10 | Tran Cris | Cris Tran | Hoang Le | LOTUS CHEMICAL TECHNOLOGY CO LTD | 147176 | 20 | CVS264944 | 4516265196 | 17623050001 | 9 | 1944 | 100% item with length out of tolerance. Spec.: 590mm +/- 10mm Actual: 520mm | Dimensional | Rework | The operator mistakenly placed SKU CVS264944 with SKU CVS792500, and the packer don't check the products before packing (IPQC). | Vendor will scrap all defect products and produce new lot with correct spec. for replacement.
283 | 2023-07-11 | Shawn Wei | Shawn Wei | Linda Zhang | HENAN HENGXIANG MEDICAL SUPPLIES CO | 120480 | 22 | 95674 | 9000845926 | 48023070014 | 30 | 600ea | table cover are not folded as drawing, non-conformance rate 100% | Visual | Rework | Not check drawing carefully | Focus on drawing and training Inspector
284 | 2023-07-11 | Shazwan | Noorul Hamsudin | Shazwan | PT JAYAMAS MEDICA INDUSTRI TBK | 153416 | 22 | 48894I | 9000846103 | 41623060002 | 44 | 8800 | 8 / 8 cases with incorrect printed PO,100% defect rate | Labeling | Rework | PPIC orders labels based on tables submitted by QC containing PO-LOT numbers and product qty. so that if there is a mistake in writing the number or qty of the QC, it cannot be detected from QC or production. | Rework by pasting corrected PO sticker on cases

## Sheet2 (hidden)
Sum of QTY | Column Labels
Row Labels | Dimension | Function | Labeling | Packaging | Visual | Visual/Material | Grand Total
rework |  |  |  |  | 1 |  | 1
Rework | 7 | 1 | 29 | 14 | 14 | 1 | 66
Rework |  |  | 1 | 3 | 1 |  | 5
Grand Total | 7 | 1 | 30 | 17 | 16 | 1 | 72
[8 empty] | Labeling | Packaging | Visual | Dimension | Function
[8 empty] | 30 | 17 | 17 | 7 | 1

## Sheet4 (hidden)
Sum of QTY | Column Labels
Row Labels | Dimension | Labeling | Packaging | Visual | Grand Total
Accept by Deviation | 34 | 15 | 17 | 9 | 75
Accept by updating PIC | 5 | 1 |  | 3 | 9
Grand Total | 39 | 16 | 17 | 12 | 84
| Dimension | Packaging | Labeling | Visual
Accept by Deviation | 39 | 17 | 16 | 12

## Reworked data (hidden)
Row Labels | Sum of QTY | Count of Defect Classification
rework | 1 | 1
Lightpak | 1 | 1
Rework | 66 | 66
FEROZE1888 MILLS | 1 | 1
Allmed | 1 | 1 |  |  |  | Dimension | Function | Labeling | Packaging | Visual | Visual/Material
Amsino | 1 | 1 |  |  | Safeway | 13 |  | 1 |  | 2 |  | 16
AYK | 2 | 2
BESMED | 2 | 2
Caina | 2 | 2 | Vendors | QTY |  | Vendors | QTY
COBES BAGO 2 | 2 | 2 | Lightpak | 1 |  | GC Medica | 7
Deming | 1 | 1 | FEROZE1888 MILLS | 1 |  | Kenpax-Anson | 5
DONGSHENG | 2 | 2 | Allmed | 1 |  | Safeway | 3
Formed | 1 | 1 | Amsino | 1 |  | Sino-Hefei | 3
GC Medica | 7 | 7 | AYK | 2 |  | SSS | 3
Gongdong | 1 | 1 | BESMED | 2 |  | Hengxiang | 3
Grand Works | 1 | 1 | Caina | 2 |  | AYK | 2
Hengxiang | 3 | 3 | COBES BAGO 2 | 2 |  | BESMED | 2
HONGDE | 2 | 2 | Deming | 1 |  | Caina | 2
Intco | 1 | 1 | DONGSHENG | 2 |  | COBES BAGO 2 | 2
Intco-Anhui | 1 | 1 | Formed | 1 |  | DONGSHENG | 2
JIEGAO | 1 | 1 | GC Medica | 7 |  | HONGDE | 2
Kenpax-Anson | 5 | 5 | Gongdong | 1 |  | LUCKY TEXTILE | 2
Longteng | 1 | 1 | Grand Works | 1 |  | RUNQIANG | 2
LUCKY TEXTILE | 2 | 2 | Hengxiang | 3 |  | Yaofeng | 2
M & G | 1 | 1 | HONGDE | 2 |  | Mediblue | 2
Meccan | 2 | 2 | Intco | 1 |  | Others | 1
Meed Skills | 1 | 1 | Intco-Anhui | 1
MF(Huangjiang) | 1 | 1 | JIEGAO | 1
Minhua(suzhou) | 1 | 1 | Kenpax-Anson | 5
Paiter | 1 | 1 | Longteng | 1
Plus Medical | 1 | 1 | LUCKY TEXTILE | 2
RUNQIANG | 2 | 2 | M & G | 1
Safeway | 3 | 3 | Meccan | 2
SESMED (CAMBODIA) HEALTHCARE | 1 | 1 | Meed Skills | 1
Sino-Hefei | 3 | 3 | MF(Huangjiang) | 1
SSS | 3 | 3 | Minhua(suzhou) | 1
Syntex | 1 | 1 | Paiter | 1
Topwide | 1 | 1 | Plus Medical | 1
WINSTAR | 1 | 1 | RUNQIANG | 2
Yaofeng | 2 | 2 | Safeway | 3
ZHENGDONG | 1 | 1 | SESMED (CAMBODIA) HEALTHCARE | 1
Rework | 5 | 5 | Sino-Hefei | 3
JIEHONG BD | 1 | 1 | SSS | 3
Mediblue | 2 | 2 | Syntex | 1
New World Imports(Accra Pac) | 1 | 1 | Topwide | 1
Seblyne-Shanghai | 1 | 1 | WINSTAR | 1
Grand Total | 72 | 72 | Yaofeng | 2
|  |  | ZHENGDONG | 1
|  |  | JIEHONG BD | 1
|  |  | Mediblue | 2
|  |  | New World Imports(Accra Pac) | 1
|  |  | Seblyne-Shanghai | 1

## Sheet3 (hidden)
Row Labels | Sum of QTY
Accept by Deviation | 75
Allmed | 1
Andon | 1
Bytech | 8
CAREMED | 2
Changlong | 2 | Vendor | QTY
COBES BAGO 2 | 1 | Safeway | 13
Deming | 1 | Bytech | 8
DONGSHENG | 1 | Topwide | 5
Eway | 1 | RR GLOBAL TRADING FZE | 4
GSPMED | 3 | Jinsheng | 4
Haoding | 1 | GSPMED | 3
JEK | 1 | Meccan | 3
Jingle | 1 | Sino-Hefei | 3
Jinsheng | 4 | Weifengrun | 3
KANGYUAN | 1 | CAREMED | 2
Lixin | 1 | Changlong | 2
Longteng | 2 | Longteng | 2
Meccan | 3 | Wellong | 2
New Berry | 1 | Wellong | 2
P&W | 1 | KDL-shanghai | 2
Rayxin | 1 | Others | 1
RR GLOBAL TRADING FZE | 4
Safeway | 13
Shen wei | 4
Sino-Hefei | 3
Sundly | 1
Topwide | 5
Weifengrun | 3
Wellong | 2
Xinbei | 1
Xinfa | 1
Accept by updating PIC | 9
Wellong | 2
AHC | 1
IMMANUEL | 1
KDL-shanghai | 2
Minhua | 1
Shuyou | 1
Viecura | 1
Grand Total | 84